Supplemental Information – Oxygen Enrichment / Nitrogen Enrichment

General

This option provides the capability to achieve and maintain the percent oxygen process values both above and below ambient levels inside Percival Scientific environmental chambers. The addition of oxygen is controlled independently, and the reduction of oxygen (oxygen flush) is achieved by the addition of nitrogen that will dilute the amount of oxygen in the air as air leaks from the chamber.

System overview

The chamber has been fitted with additive O2 and N2 systems. They have been set up in a closed-loop control scheme. Once a desired O2 level has been input into the Intellus controller, the controller will activate the appropriate system to add O2 if needed to raise the level, or add N2  to reduce the O2 level, until the input value is achieved. This is done through a feedback system where the controller measures the current O2 level inside the chamber and compares it to the input value. This loop continues until a steady state is reached.

Installation

## Gas Supply Requirements

  - Customer supplied O2 tanks should be regulated to 8-10 psig and have sufficient capacity to reach the set point and maintain required supply

  - Customer supplied N2 tanks should be regulated to 8-10 psig and have sufficient capacity to reach the set point and maintain required

  - Supply lines with appropriate fittings for your chamber; either a push-type for a ¼” o.d. tube or a ¼” Male SAE

Connections

The O2 and N2 systems are equipped with needle valves to connect the lines

- Connect the appropriate gas supply line to the provided needle valves

- The valves controlling the flow of O2 and N2 are set at the factory and should not need adjustment. The ideal pressures supplied to the system are subject to chamber size, experiment type, and desired system response and controllability.

- Should the need to adjust the valve be necessary.

  - Remove the clip ring

  - Slide the red ring up and turn the dial to adjust flow

  - Adjust flow to desired setting

  - Slide red ring down and reinsert clip.

Connect gas supply here (SAE fitting shown)

Red ring shown in locked position

Install clip to lock valve position

The pressure on the down-side of the regulator should not be set higher than 20psig. A reasonable supply pressure is approximately 8-10psig.

The ideal N2 pressure supplied to the system is subject to chamber size, experiment type and desired system response and controllability. Keep in mind that raising the supply pressure to the N2 solenoid valve will increase the system's responsiveness but will decrease overall controllability. Alternatively, lowering the N2 supply pressure will have the opposite effect on the system.

System maintenance

- Periodically check the O2 and N2 supply tank pressures and replace the tanks when the supply pressure gets too low.

- Periodically make sure all fittings in the system are tight.

Control

The oxygen enrichment system is controlled via channel A. The nitrogen enrichment is controlled via channel B. There is a safety system that runs separately and prevents the oxygen percentage from exceeding the OSHA limit of 23.5%. Please refer to the attached Intellus Controller Manual for information on using the controller to create programs or run manual settings.

Service Information

Disconnect all unnecessary electrical power when servicing equipment.

Majority of circuits are high voltage.

- Intellus controller: Check to ensure that “Enable Aux Output 4” is set to “Yes” in the Intellus controller when the additive N2 system is being utilized. If additive N2 is not necessary, set “Enable Aux Output 4” to “No”.

- Intellus controller: Check the controller for correct voltage output to the additive N2 relay. If the controller fails to output the correct voltage (approximately 5VDC), replace the controller.

- Relays: Check for proper function of relays.

- O2 sensor: If the reading on Intellus controller is incorrect, check to ensure that O2 sensor is functioning properly.

- Needle valves and solenoid valves: Check to ensure they are open and check for proper function of solenoid valves.

- Supply tanks: Check to ensure supply tanks are not too low on pressure. If it is, replace tank.

Recommended maintenance schedule: At customer’s discretion.