AL-30L2 Chamber

Operator’s Manual

505 Research Drive  Perry, Iowa 50220

Toll-free 800.695.2743  515.465.9363  fax 515.465.9464

www.percival-scientific.com

[This page intentionally left blank]

# Introduction

To help customers create better science, Percival Scientific continuously studies and refines the design of its products to provide the best possible solutions and equipment in the industry.  As such, changes in the design and/or operation of equipment may occur without publication.  Please feel free to contact Percival Scientific directly for product updates.

Please read this manual carefully before proceeding with operation

# Revision Log

| Version | Date | Description |
| --- | --- | --- |
| Rev A | 4/18/25 | Document released from draft |

Introduction	3

Revision Log	4

Warranty Information	7

Definitions	9

General	9

Equipment Overview	9

Product Specifications	10

Installation	11

Site Preparation	11

Location	11

Ambient Conditions	13

Chamber Heat Rejection to Ambient	13

Electrical Services	13

Drainage and Plumbing Services	15

Unpackaging	16

Inspect the Shipment	16

Remove the Chamber from its Skid	17

Moving the Chamber	17

Installation Guidelines	17

Level the Unit	17

Inspect the Control Box	18

Install Shelving and Lights	18

Connect Services	20

Startup and Commissioning	21

Turn on the Chamber	21

IntellusUltra Control System	21

Calibration	22

Operation	24

Guidelines for Loading a Chamber	24

Running a Control Profile	24

Air Flow	25

Perforated Plenums	25

Maintenance	26

General Maintenance Summary Table	27

Cleaning the Chamber	27

Clean the Heat Exchangers	27

Clean the Condenser	27

Clean the Evaporator	28

Troubleshooting	31

Loss of Temperature Control – General	31

Loss of Temperature Control - Overheating	31

Loss of Temperature Control – Overcooling	31

General	32

Parts and Service	32

# Warranty Information

Equipment Limited Warranty

Thank you for purchasing equipment from Percival Scientific, Inc.  Our goal is to provide every customer with an exceptional ownership experience.  We are committed to providing products and services of the highest quality, and your satisfaction is our top priority.  In support of this objective, we present the following warranty.

Warranty Terms

Percival Scientific, Inc. (hereinafter referred to as Percival) warrants that equipment manufactured by Percival shall be free from defect in material or workmanship, when installed, operated and maintained in accordance with instructions provided by Percival.

Warranty Coverage

- Coverage period shall be twenty-four (24) months from the date of delivery

- Percival’s obligation under this warranty is limited to the repair or replacement of defect at no charge for parts, shipping and labor

- Percival reserves the right to inspect the equipment after a warranty claim is made to determine the proper course of corrective action

Exclusions

- Parts or elements damaged from improper installation, including mechanical, electrical and plumbing services, IT infrastructure, inadequate room ventilation, improper water quality, and ambient conditions that are out of specification

- Parts or elements damaged from improper operation

- Parts or elements damaged from improper maintenance

- Parts or elements considered to be consumable, that outlive their normal expected terms of usefulness, such as: light bulbs, filters, desiccant material, and CO2 scrubbing media)

- Damage resulting from shipment to or relocation within the customer’s facility after equipment has been signed for and ownership is transferred to the customer

- Modification to equipment by the customer without Percival’s consent may void this warranty

- Repair work undertaken without prior consent of Percival may void this warranty

- Damage caused by accident, fire, Force Majeure or Acts of God

Additional Terms

- This warranty is in lieu of any other warranties, expressed or implied, including merchantability or fitness for a particular purpose

- Warranty terms described herein cannot be changed without the expressed, written consent from Percival

- The owner agrees that Percival’s sole liability with respect to defective parts or materials shall be set forth in this warranty, and any claims for incidental or consequential damages are expressly excluded

- This warranty is non-transferable and applies only to the original purchaser

- Repaired or replaced parts shall be warranted only for the unexpired portion of the original warranty, unless otherwise defined

- Parts replaced under warranty shall become the property of Percival

To Make a Warranty Claim

- Document the issue with detailed notes and photographs along with the equipment’s serial number

- Contact Percival’s customer support team at customerservice@percival-scientific.com, parts@percival-scientific.com or 1-800-695-2743

- Once your claim is received, our customer support team will review it and contact you with further instructions

# Definitions

‘Controlled environment’, ‘Chamber’ and ‘Unit’ are used interchangeably in this manual to describe the complete controlled environment chamber system manufactured by Percival Scientific, Inc.

# General

This manual is intended to provide instructions on the installation, operation, maintenance and troubleshooting of Percival Scientific’s AL-30L2 Algae growth chamber.

Algae chambers play a crucial role in various scientific and medical fields, serving as controlled environments for the growth and study of algal samples.  Common research applications include biodiesel fuels, pollution control, bacterial growth mediums, carbon dioxide fixation, photosynthesis, food additives, and coral-algal symbiosis.

## Equipment Overview

## Product Specifications

| Model |  |  | AL-30L2 |
| --- | --- | --- | --- |
| Temperature Range | Lights Off | Low | 2.0°C |
|  |  | High | 44.0°C |
|  |  | Range | ±0.5°C |
|  | Lights On | Low | 7.0°C |
|  |  | High | 44.0°C |
|  |  | Range | ±0.5°C |
| Exterior Dimensions |  | Width | 31.0 in (78.7 cm) |
|  |  | Depth | 24.4 in (62.0 cm) |
|  |  | Height | 46.2 in (117.3 cm) |
| Interior Volume |  | ft³ | 9.86 |
|  |  | m³ | 0.279 |
|  |  | L | 279.2 |
| Number of Tiers |  |  | 2 |
| Shelves per Tier |  |  | 1 |
| Shelf Dimensions |  | Width | 26.5 in (67.3 cm) |
|  |  | Depth | 16.5 in (41.9 cm) |
| Total Shelf Area |  | ft² (m²) | 6.1 (0.56) |
| Growing Height per Tier |  | in (cm) | 10.4 (26.4) |
| Door Opening |  | Width | 26.7 n (67.8 cm) |
|  |  | Height | 29.8 in (75.7 cm) |
| Light Intensity @ 6” (µmol/m²/s) |  |  | 440 |
| Electrical |  | VAC | 120V |
|  |  | Phase | 1 |
|  |  | Frequency | 60Hz |
|  |  | Cords (NEMA) | 1 (5-15P) |
|  |  | RLA | 9.9A |
|  |  | MCA | 12.4A |
| Heat Rejection to Ambient |  | Btu/h (W) | 2054 (602) |
| Shipping Weight |  | lb (kg) | 328 (148.8) |

# Installation

To achieve and maintain optimum chamber performance throughout its designed operating range(s), it is important to install the unit according to the following guidelines.  Failure to follow the installation guidelines in this section may cause loss of environmental control, malfunction, premature system failure and may also void the manufacturer’s warranty.

## Site Preparation

The following considerations should be made when preparing a facility for installation of Percival controlled environments:

### Location

Locate the chamber in a clean area with a level surface and adequate space for door swings, ventilation, access of facility services, and the unit’s service connections and access panels.

Prior to the unit’s delivery date, ensure there is a clear path from the delivery dock to final chamber location, taking doorways, elevators, corridor widths and turns into account.

Warning! The AL-30L2 chamber is heavy.  Please use caution when removing the unit from its shipping container and while moving it into position.  These tasks should be undertaken with care by more than one person.

Spatial requirements vary by chamber model and optional equipment.  Please refer to all supplementary installation instructions as optional equipment can change the space required for a chamber’s proper operation.

Provide at least six inches of clearance on the right side and behind the AL-30L2 unit to facilitate proper air flow through the condensing unit’s heat exchanger.  Air is drawn in through the right-handed side vent and exhausted through the back vent of the unit’s mechanical section.

Also, ensure there is sufficient space in front of the chamber to allow the door to open at least 90 degrees for full access to the chamber’s interior.  The door can open as far as 180 degrees and additional clearance for opening the door beyond 90 may allow easier access for working in the chamber.

Minimum space requirements for installation of the standard AL-30L2 controlled environment chamber are shown in the footprint drawing below.

### Ambient Conditions

Published chamber performance specifications are based on the following ambient environmental conditions.  To ensure best performance of the AL-30L2 controlled environment chamber, locate the unit in an environment where the following conditions can be maintained:

Temperature = 75 °F (23.8 °C)

Relative Humidity = 55 %RH

### Chamber Heat Rejection to Ambient

To achieve and maintain specified conditions within a controlled environment chamber, the unit’s cooling system must remove excess heat from within the chamber.  This excess heat is rejected from the system and must be accounted for when planning for a chamber’s installation.

The base AL-30L2 chamber model’s estimated maximum heat rejection to ambient can be found in the Product Specifications section of this manual.

It is necessary to ensure the HVAC system where the chamber is installed has sufficient capacity to remove heat rejected by all chambers within the space.  Insufficient heat removal capacity will result in elevated ambient temperatures and possible degradation in chamber performance.

### Electrical Services

The AL-30L2 controlled environment chamber is typically provided with one or more power cords that must be plugged into electrical building services of appropriate type, rating and ampacity.

Always verify a chamber’s electrical requirements against the provided electrical services prior to making final connections and energizing the unit.

A chamber’s electrical requirements are listed on its serial tag, which is located inside the unit’s door on the interior side wall of the controlled environment space.  These electrical requirements include the operating voltage, phase(s) and maximum running load amperage (RLA) for the chamber, per the following requirements:

- Supply voltage must be within ±10% of the serial tag value.

- The number of phases provided should match the serial tag value.

- The electrical service must have a minimum circuit ampacity (MCA) sufficient to support the chamber’s rated maximum running load amperage (RLA).  Service MCA should ultimately be determined by local authorities having jurisdiction (AHJ), according to state, local and building electrical codes.  As a rule of thumb, MCA can be estimated by the following equation:

MCA=1.25×RLAmax

- The provided electrical service should not be smaller than the MCA.  In many cases the service itself will be larger than the MCA. For example, a standard 30A electrical service might be specified to comply with a chamber’s MCA requirement of 27A.

If a chamber is provided with multiple power cords, each cord will have its own RLA/MCA requirement.  If both cords are to be connected on a single circuit with multiple receptacles, the service must be sized for the COMBINED RLA of both cords.  An example equation is provided below for reference:

MCAcombined=1.25 ×(RLAcord 1+ RLAcord 2)

In lieu of a single electrical service sized for multiple cords, connect each power cord to its own separate, dedicated electrical service.

Failure to follow the electrical service requirements outlined above can result in overloaded electrical services, tripped circuit protection on the supply and damage to facility and chamber electrical infrastructure such as wiring, plugs and receptacles.

Damage caused from incorrect or improperly sized electrical services is not covered under the manufacturer’s warranty.

The duplex receptacle shown above can physically accommodate two 20A rated cords but may not have sufficient capacity to supply power at the full rating of two cords, simultaneously.  Verify the combined load the duplex receptacle’s electrical service can safely support before plugging in multiple power cords.

Do not plug multiple chamber power cords into a duplex receptacle unless the circuit is dedicated and verified to have an MCA sufficient to support the combined RLA of both connected power cords.

Check quadplex receptacles for proper labelling indicating the presence of multiple circuits prior to plugging in multiple chamber power cords

### Drainage and Plumbing Services

The model AL-30L2 controlled environment rejects condensed water from its evaporator coils.  The amount of condensate is dependent upon a combination of controlled environment settings, ambient conditions, door opening frequency and duration, air exchange through (optional) fresh air ports, and characteristics of products placed inside the environment.

Optional relative humidity control systems increase the amount of condensate rejected from a given unit.

Whenever the evaporator coil operates at a temperature below the dew point inside the controlled environment, moisture present in the circulated air condenses onto the cooling coils.  This condensate sheds off the coils, is routed through a drain and expelled through the environment’s floor drain and out the back of the unit’s mechanical compartment.

Base model AL-30L2 algae units are provided with one 7/8 in ID (1 in OD) flexible drain line.  Ideally, this drain line should be routed toward a floor drain close in proximity to the chamber to properly manage condensate rejected from the unit.

If the installation location for the chamber does not provide a nearby floor drain, the following drain accommodations can be considered:

- Extend the existing chamber drain tubing and route it to a floor drain that is farther away than the chamber’s footprint.

- Place the existing drain tubing into a Q19 (External Drip Pan) optional equipment package.

- Connect the existing chamber drain system to a condensate pump

Please note that the standard drain system in an AL-30L2 controlled environment chamber works via gravity.  Condensate will only drain properly if the facility drain is lower in elevation than the chamber floor drain and tubing.

## Unpackaging

### Inspect the Shipment

At the time of delivery, it is very important to verify that the proper quantity and type of products were received as per the Bill of Lading, and to inspect those products for shipping damage prior to accepting the shipment.

- Carefully remove all packaging materials that surround the unit, ensuring no loose chamber components are accidentally discarded.

- Visually inspect both the interior and exterior of the unit for dents, scratches or any other signs of damage

- It is extremely important to note any damage on the shipping receipt so that corrective actions can be taken as efficiently and expeditiously as possible.  The shipping documentation provides the following options:

  - Accept the shipment as-is

  - Accept the shipment with damages documented, understanding that corrective actions to rectify the damage will be performed onsite

  - Reject the shipment, where damage is sufficient to require the product be returned to Percival, where corrective actions can be taken at the manufacturing facility

### Remove the Chamber from its Skid

Warning! The AL-30L2 chamber is heavy.  Please use caution when removing the unit from its shipping container and while moving it into position.  These tasks should be undertaken with care by more than one person.

- Remove shelving and any additional items packaged inside the chamber that can shift when the chamber is moved

- If there is a drain line affixed to the shipping skid, free it and temporarily secure it to the chamber to prevent it from hindering movement of unit

- Remove any metal shipping brackets that affix the chamber to the skid

- Remove any bolts that affix the chamber to the skid

- Lift the chamber off its shipping skid and place it upright onto a level surface

### Moving the Chamber

Do not attempt to move any AL-30L2 chamber loaded with experiments, samples or other sensitive materials.

Warning! The AL-30L2 chamber is heavy.  Please use caution while moving the chamber.  This task should be undertaken with care by more than one person.

Warning! Once removed from it skid, sliding the AL-30L2 into position may scratch the surface upon which the unit is placed.

## Installation Guidelines

Once placed in a suitable location as prescribed in the preceding sections of this user manual, proceed with the following installation steps:

### Level the Unit

Make sure the unit is placed upon a level surface to facilitate proper function and drainage.  The base model AL-30L2 chamber is not provided with a means for levelling the unit on uneven or sloped surfaces.  Levelling legs and casters are available as optional equipment that can be retrofitted onto the AL-30L2 chamber base.  Please contact Percival’s Customer Support team at customerservice@percival-scientific.com, or 1-800-695-2743.

### Inspect the Control Box

While unplugged from electrical services, open the chamber’s control panel and perform a visual inspection of the electrical components installed within it.

- Check the state of the resettable circuit breakers and reset any found in an ‘open’ state

- Check for any components that are not physically secured to the component mounting plate, control box housing or top of the insulated box.

- Check for any connector plugs that are dislodged or not fully seated into their receptacles and re-secure them as needed with a firm press.

- Check for any obviously loose or disconnected wiring

If any components have become dislodged or there is any question as to where a disconnected plug or wire should be connected, do not proceed any further with chamber installation or operation and contact Percival’s customer support team at 1-800-695-2743 or customerservice@percival-scientific.com

### Install Shelving and Lights

The AL-30L2 controlled environment chamber is provided with two (2) steel wire shelves and two (2) SciWhite® lamp banks, which are shipped uninstalled and packaged within each chamber.  For additional information regarding the SciWhite® lighting system please refer to the SW - SciWhite® Series LED Lighting Operator’s Manual.

To install shelving and lamp banks in an AL-30L2 base model chamber:

- Unpackage each shelf and lamp bank and locate the pilaster clips

- Install pilaster clips:

  - Four clips, located at the same height are required to properly support each shelf and lamp bank

  - Numbers are embossed into each pilaster to assist with alignment of the clips

  - Insert the upper pilaster clip tab into the pilaster first

  - Twist the pilaster clip in a downward motion and insert the lower tab into the lower pilaster slot.

Note:  It may be necessary to pinch the clip slightly, upon insertion, to obtain proper engagement of both clip tabs into the pilaster.

The image below shows an example of a properly installed pilaster clip:

- Where to locate the clips:

  - Place four clips so that the upper lamp bank can be placed in the highest position near the top of the pilasters and four clips near the bottom for the lower shelf

  - Bisect the remaining space with the upper shelf directly above the lower lamp bank (see image on left below)

  - When evenly spaced, each tier will provide approximately 10.4-inch growth height.

As lamp banks and shelves are adjustable, uneven spacing of each tier is possible.  Uneven spacing can lead to differences in the chamber’s specified air flow, control tolerances, spatial uniformity and light intensity.

- Place shelves and lamp banks onto the installed pilaster clips, ensuring that the vertical tab on each clip sits inboard of their frame, which secures each in place.  When properly installed, a shelf or lamp bank can only be removed by lifting it vertically off its clips. (see image on left below)

(Tabs on pilaster clips must sit inside shelf frame and lamp bank flanges)

- Make sure to orient each lamp bank assembly such that power cords extend toward the back wall of the controlled environment or are close in proximity to the chamber’s corresponding receptacles.

Ensure chamber power is turned off, prior to plugging in lamp banks

- Insert the male power plug of each SciWhite® Series lamp bank into its corresponding female receptacle.  The plug and receptacle are configured such that they can only be connected in one orientation.

Make sure each plug is fully seated into its receptacle with a firm press.  Failure to fully seat a plug into its receptacle may cause the lamp bank or chamber electrical system to malfunction or become damaged.

(Plug not fully seated)		             (Press plug firmly into receptacle)

### Connect Services

- Plug the unit into properly sized electrical services, ensuring that all plugs are completely and securely inserted

- Route the chamber’s drains to a floor drain or other drainage accommodation

- Connect services for optional equipment as directed in supplemental operation manuals (e.g. demineralized water, compressed gases, chilled water, ventilation ductwork, etc...)

## Startup and Commissioning

### Turn on the Chamber

Set the main power switch located on the left side of the control panel to the ‘On’ or ‘|’ position to energize the chamber.  The IntellusUltra controller takes approximately 15-20 seconds to start up.  Once the controller has started up, and there are no active alarm conditions, the rest of the chamber components will energize. Depending on the options and chamber setup, the full start up process can take several minutes.

### IntellusUltra Control System

The primary method of operating an AL-30L2 controlled environment chamber is through the C8 IntellusUltra control system’s touch screen. The touch screen (shown below) is located on the exterior of the AL-30L2 Series base model’s control panel.

Please refer to the supplementary Touch Screen User Manual for instructions on the operation of this system.

Refer to the IntellusUltra Quick Guide provided with the chamber for basic instructions on the controller’s operation.  This guide can also be downloaded at https://www.percival-scientific.com/wp-content/uploads/2023/03/IntellusUltra-Quick-Guide.pdf

Refer to the C8 Control System User Manual for comprehensive operational instructions for the chamber’s control system

When first energized, the IntellusUltra control system is set to operate in Manual Mode, and has the following factory settings:

| Control Feature | Factory Setting |
| --- | --- |
| System Time | Current North American Central Time Zone |
| Operation Mode | Manual |
| Programming Step Type | Non-ramping |
| Temperature Set Point | 25°C |
| Temperature Limit 1 High Alarm | 35°C |
| Temperature Limit 1 Low Alarm | -5°C |
| Temperature Limit 2 High Alarm | 38°C |
| Temperature Limit 2 Low Alarm | -8°C |
| Light Output 1 | 0% (Off) |

In general, any optional features controlled via the IntellusUltra are disabled by default.  Please refer to the user manuals provided with optional equipment for information pertaining to their factory settings, setup and operation.

Note: The keypad overlay (shown below) inside the control box is used for service purposes.

### Calibration

All controlled environment chambers from Percival Scientific must pass rigorous quality checks and validation testing prior to shipment.  As part of the validation process, all chamber control systems are calibrated against an NIST standard device at the following conditions:

| Condition | Calibration |
| --- | --- |
| Temperature w/ Lights Off | 25°C |
| Temperature w/ Lights On (for lighted models) | 25°C |
| Relative Humidity w/ Lights Off (if RH option is equipped) | 75%RH @ 25°C |
| Relative Humidity w/ Lights On (if RH option is equipped) | 75%RH @ 25°C |

At these calibration points, any discrepancy between an IntellusUltra sensor (located by the intake of the evaporator coil) reading and the NIST calibration standard (located in the center of the chamber environment) are entered into the control system as calibration offsets.

For any given environmental parameter controlled by the chamber:

Calibration Offset=Calibration Standard Value-Control System Process Value

For example, if the control system reads a temperature process value of 25°C and the calibration standard reads 24.5°C, then an offset of -0.5°C is saved into the controller.  For each condition, the measured offset is applied to the process value so that it matches the calibration standard.

TOffset= TStandard- TProcess Value

TOffset= 24.5°C- 25.0°C

TOffset= -0.5°C

Any offsets with lights turned off are called Night Offsets and any offsets with any number of lights turned on are called Day Offsets.

Please refer to the provided control system user manuals for additional information regarding calibration offsets.

Factory calibration offsets serve as suitable values for general chamber use.  Consider recalibrating the chamber if precise environmental control is required at specific conditions that are significantly different than the factory calibration points.

# Operation

## Guidelines for Loading a Chamber

When placing products, specimens or experiments onto shelves inside an AL-30L2 controlled environment, adhere to the following guidelines to ensure the unit operates at peak performance:

- For best temperature uniformity, evenly distribute items throughout the chamber space and allow space between items to promote proper air flow

- As a good rule of thumb, do not populate more than 80% of any shelf’s surface area

- Do not place more than 125 lbs (56.8 kg) onto any shelf

- Do not place more than 500 lbs (227.3 kg) total in a chamber model

Avoid placing items in a unit that produce significant amounts of heat.  AL-30L2 base model chambers are designed to accommodate a product heat load of 3 W/ft² of shelf space.  Any heat introduced into the chamber environment beyond this amount can result in loss of environmental control.

## Running a Control Profile

- Close the main chamber door, ensuring a proper seal between the door gasket and insulated box

- Configure static or multi-step profile process value set points

- Check alarm limits, ensuring that they are set above the maximum and below the minimum process values.  Limits should not be set closer than 2-3°C above or below the operational range of the entire control profile.

- Select the mode of operation on the control system (e.g. Manual, Diurnal, Program or Sequence).

Once an operating mode is set, the IntellusUltra control system automatically executes any set points specified in the profile.

If an alarm threshold is exceeded, the controller produces an audible alarm until the alarm is acknowledged by the user.  The alarm state remains active until the source of the alarm is resolved.  A temperature alarm condition shuts down all chamber sub-systems inside the environment until the alarm state is cleared.  Whenever process values fall back within allowable limits, the control system automatically reverts to normal operation, requiring no involvement from the user.

The IntellusUltra controller has an internal battery which allows it to keep time in the event of a power cycle or outage.  Upon power up, the chamber will resume the previously executed control profile.

- Control is resumed around static process value set points when the controller is in manual mode

- When executing a multi-step program, the controller resumes control around the current program step as determined by the controller system time.

- When executing a sequence of programs, the controller resumes control around the current program iteration and step as determined by the system time unless power is off at 12:00 am.  If power is off at 12:00am the program iteration count will not be incremented.

For best control, minimize the frequency and duration of chamber door openings.

## Air Flow

Air flow across each tier in the AL-30L2 is horizontal.

The chamber’s temperature control system continuously circulates air in the following manner:

- Air is drawn into the ceiling-mounted evaporator coil for conditioning

- Conditioned air cascades down the rear wall of the chamber and pressurizes behind a perforated plenum

- The perforated plenums promote uniform air flow across each tier, where process air flows horizontally from back to front of the experiment space

- Air returns upward at the front of the chamber, between the shelves and door, where it is drawn back into the evaporator coil for reconditioning

Air flow is tuned to provide sufficient capacity to remove excess heat from within the chamber and establish proper uniformity while minimizing disturbance to specimens.

## Perforated Plenums

Perforated rear wall plenums are used to develop uniform, horizontal air flow across each tier.  The plenum affords some adjustment of the air flow amount across each tier.  This is accomplished through an adjustable panel that can be used to restrict flow through each plenum section.

Temperature differences between shelf tiers can be minimized through adjustment of the plenums.  As a rule of thumb, adjusting a plenum section open will increase the rate of air flow, and decrease the shelf temperature relative to the other tiers.  Adjusting a plenum closed will have the reverse effect.  When adjusting plenums, please keep in mind that there is a fixed total amount of air movement capacity in the chamber system. Adjustment of one plenum section will either steal air from or divert excess air toward the other plenum sections.

Do not remove the screws or the air block panel will detach.

To adjust a plenum section, loosen the adjustment screws only enough to allow the air block panel to slide horizontally.  Slide a plenum adjustment panel left to increase air flow and right to reduce air flow.  Once the panel has been adjusted, retighten the screws.

Air Block

Adjustment Screw

Air Block

Adjustment Screw

# Maintenance

Performing regular maintenance on an environmental chamber reduces performance degradation, improves reliability and extends operational life.  In addition to any user-specific maintenance protocols, the following general maintenance tasks should be performed at the recommended intervals.

Turn chamber power off prior to performing any maintenance inside the environment.

Unplug or disconnect chamber supply power prior to performing any maintenance inside the chamber’s mechanical compartment or control box.

## General Maintenance Summary Table

| Maintenance Task | Recommended Interval Schedule |
| --- | --- |
| Clean the Chamber | Every 3 months, or as needed |
| Clean the Condenser | Every 3 months, or as needed |
| Clean the Evaporator | Every 12 months |

## Cleaning the Chamber

Wipe down all interior and exterior surfaces with a solution of water and mild detergent.  Do not use abrasive agents or cleaning pads as these could scratch surfaces.

Glass cleaners can be used on stainless steel surfaces to prevent streaking.

## Clean the Heat Exchangers

Air must pass freely through the evaporator and condenser for proper performance.  Ambient air is drawn into and forced through the condenser by the condensing unit fan.  Chamber process air is drawn into and forced through the evaporator by a bank of fans inside the evaporator cover.  In both cases, the fans run continuously while the chamber is in operation and there are no active alarms.

Verify the chamber power is off prior to any cleaning effort on refrigeration system components.

Be aware that some parts of the condensing unit become hot while the unit is in operation and can remain hot for some time after the unit is powered down.

Heater elements located behind the evaporator coil may become hot during operation and should be allowed to sufficiently cool before any maintenance activities on the evaporator coil are undertaken.

### Clean the Condenser

Once powered down, remove any dust or debris collected upon the condenser fins, copper tubing, fan blades and motor.  Blow dust out of the condenser from behind with compressed air, a brush, or vacuum and wipe dust from the fan blades.

To access the condenser coil, remove the vented side cover panels on the unit’s mechanical section.

To access the condenser fan, remove the vented side and rear cover panels on the unit’s mechanical section.  If necessary, the chamber’s control box can be removed from the mechanical section.  Remove the machine screws securing the control box in place and slide the unit away from the mechanical section.  There is sufficient wiring to allow access to mechanical components from the front of the unit without the need to disconnect any wire looms. Special care should be taken through the disassembly, cleaning, and the reassembly process to not damage any wires.

(Remove inner machine screws and top hinge screws for control box removal)

### Clean the Evaporator

The following steps and images are for reference on how to remove the evaporator cover to access the evaporator coil and air circulation fans located in the cover for cleaning.

- Verify that the chamber power is off prior to cleaning the evaporator components

- Unplug any lights or auxiliary equipment plugged into the chamber’s receptacle panel that will prevent access to the evaporator

- Remove lamp banks and shelving that may interfere with removing the cover. Pilaster clips can often be left in place

- Remove the two outside sheet metal screws on the left that secure the evaporator cover to its housing and gently let it hang to access to the fan wiring

- Slide the clip with the ground (green) wire off the frame and unwrap it from around the other wires

- Disconnect the neutral (white) wire from the ribbed wire that runs to the fans and the positive (black) wire from the smooth wire that also runs to the fans

Note: Chambers with electrical requirements other than 115/1/60 may have two black wires and no white. In this case the wires should be marked before disconnecting

- Remove the other two screws holding the cover in place and set the cover with fans and screws to the side

- Clean the evaporator coil by dislodging dirt and debris via soft brush, compressed air or vacuum, and inspect the condensate drain line for any cracks, wear or blockages.

- Remove dust and debris from the evaporator fans and wipe any dirt or residue from inside the evaporator cover and housing

- Reverse the steps to replace the cover, making sure to reconnect the wires as they were originally and wrap the ground wire around the others to keep it from getting into the fans

# Troubleshooting

## Loss of Temperature Control – General

| Possible Reason | Solution |
| --- | --- |
| Improper or incorrect signal between controller and control relay(s) | Check for ~5VDC output signal between the controller and solid-state relay. Output state is indicated by LEDs on the controller overlay. Contact factory if insufficient output voltage from the control system |
| Temperature sensor | Check the sensor for correct output signal. Refer to the RTD temperature vs. resistance table in the controller manual. Contact factory if sensor output does not match the table values. |
| Solid-state relays | If correct trigger voltage is correct to the relay coil, check for the correct output voltage through the relay. Refer to the chamber’s electrical schematic for the VAC output to the affected heating and/or cooling device(s) |
| Hot gas solenoid valve | When in heating mode, the hot gas solenoid should be in the open position allowing hot gas to flow into the evaporator coil. Check the solenoid valve for proper actuation and the valve body for internal blockages. |
| Liquid line solenoid valve | When in cooling mode, the liquid line solenoid valve should be in the open position, allowing liquid refrigerant to flow into the thermal expansion device(s). Check the solenoid for proper actuation and the valve body for internal blockages. |

## Loss of Temperature Control - Overheating

| Loss of refrigerant | Check system pressures and check for leaks with a leak detector. Evacuate the system, repair leak(s) and recharge the system. Refer to the unit’s refrigeration diagram and serial label. |
| --- | --- |
| Restriction in the refrigeration piping or componentry | If adequate pressure and no leaks, there may be a restriction in the unit’s liquid line solenoid valve, filter/drier or thermal expansion device. Check the line temperature directly before and after each device. There may be a blockage if the outlet is colder than the inlet to the solenoid and drier. |
| Insufficient heat exchange through the condenser | Check the ambient air temperature is below 90°F. Make sure the condenser fan and coils are clean. Make sure the condenser fan is running. |
| Compressor motor fails to pump refrigerant throughout the system | Check electrical power to the motor and start components. Perform an ohm check on the motor windings. Check high-side system pressure if the unit runs |
| Insufficient heat exchange through the evaporator | Check for proper fan operation. Fans should be continuously running, drawing air into the evaporator. Make sure the evaporator coils are clean. |

## Loss of Temperature Control – Overcooling

| Electric heater malfunction | Check output voltage from the controller to the heater relay. Check output voltage from the heater relay to the heater element. Check the circuit breaker in the heater electrical circuit. |
| --- | --- |
| Malfunctioning hot gas solid-state relay (capillary tube systems) | Check output voltage from the controller to the hot gas line solenoid valve. Check voltage between the relay and solenoid valve. The solenoid valve should be open whenever the ‘heat’ LED is lit |
| Malfunctioning liquid solid-state relay (TXV systems) | Check output voltage between controller and relay. Check voltage between relay and solenoid. The solenoid should fully close when the ‘heat’ LED indicator light is energized. |
| Blocked hot gas solenoid valve (capillary tube systems) | If the hot gas solenoid valve gets blocked, then flow will divert through the capillary tube, causing the system to cool even when calling for heat. |

## General

| Chamber does not power on | Check facility electrical services and circuit breakers. Check that the chamber is properly plugged into the facility services. Check circuit breakers inside the unit’s control box. |
| --- | --- |
| Individual chamber component does not turn on | Check controller settings and verify that the device has been properly enabled. Check the control box for tripped circuit breakers. |
| Condensation | The amount of condensation that can form inside and be rejected from a controlled environment can fluctuate, and is highly dependent upon ambient conditions, interior environmental conditions, fresh air exchange through (optional) fresh air ports, frequency of door openings, product load and optional humidity control systems. Minimize door openings and the amount of moisture present inside the unit, and close the access port |

## Parts and Service

Replacement parts can be ordered from Percival Scientific at:

Ph: 1-800-695-2743

Email: parts@percival-scientific.com

For additional questions regarding this product, please contact Percival’s customer support team at:

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com