I-Series Chamber

Operator’s Manual

505 Research Drive  Perry, Iowa 50220

Toll-free 800.695.2743  515.465.9363  fax 515.465.9464

www.percival-scientific.com

[This page intentionally left blank]

# Introduction

To help customers create better science, Percival Scientific continuously studies and refines the design of its products to provide the best possible solutions and equipment in the industry.  As such, changes in the design and/or operation of equipment may occur without publication.  Please feel free to contact Percival Scientific directly for product updates.

Please read this manual carefully before proceeding with operation

# Revision Log

| Version | Date | Description |
| --- | --- | --- |
| Rev A | 10/7/2024 | Initial Release |

Introduction	3

Revision Log	4

Warranty Information	7

Definitions	9

General	9

Equipment Overview:	10

Product Specifications	14

I22 Series Specification Table	14

I36 Series Specification Table	15

I41 Series Specification Table	16

I66 Series Specification Table	17

Installation	18

Site Preparation	18

Location	18

Ambient Conditions	19

Chamber Heat Rejection to Ambient	19

Electrical Services	19

Drainage and Plumbing Services	21

Unpackaging	22

Inspect the Shipment	22

Remove the Chamber from its Skid	23

Moving the Chamber	23

Installation Guidelines	23

Level the Unit	24

Inspect the Control Box	24

Install Shelving and Lights	25

Connect Services	26

Startup and Commissioning	27

Turn on the Chamber	27

IntellusUltra Control System	27

Calibration	28

Operation	30

Guidelines for Loading a Chamber	30

Running a Control Profile	30

Maintenance	32

General Maintenance Summary Table	32

Cleaning the Chamber	32

Clean the Heat Exchangers	32

Clean the Condenser	33

Clean the Evaporator	34

Troubleshooting	35

Loss of Temperature Control – General	35

Loss of Temperature Control - Overheating	35

Loss of Temperature Control – Overcooling	36

General	36

Parts and Service	37

# Warranty Information

Equipment Limited Warranty

Thank you for purchasing equipment from Percival Scientific, Inc.  Our goal is to provide every customer with an exceptional ownership experience.  We are committed to providing products and services of the highest quality, and your satisfaction is our top priority.  In support of this objective, we present the following warranty.

Warranty Terms

Percival Scientific, Inc. (hereinafter referred to as Percival) warrants that equipment manufactured by Percival shall be free from defect in material or workmanship, when installed, operated and maintained in accordance with instructions provided by Percival.

Warranty Coverage

- Coverage period shall be twenty-four (24) months from the date of delivery

- Percival’s obligation under this warranty is limited to the repair or replacement of defect at no charge for parts, shipping and labor

- Percival reserves the right to inspect the equipment after a warranty claim is made to determine the proper course of corrective action

Exclusions

- Parts or elements damaged from improper installation, including mechanical, electrical and plumbing services, IT infrastructure, inadequate room ventilation, improper water quality, and ambient conditions that are out of specification

- Parts or elements damaged from improper operation

- Parts or elements damaged from improper maintenance

- Parts or elements considered to be consumable, that outlive their normal expected terms of usefulness, such as: light bulbs, filters, desiccant material and CO2 scrubbing media)

- Damage resulting from shipment to or relocation within the customer’s facility after equipment has been signed for and ownership is transferred to the customer

- Modification to equipment by the customer without Percival’s consent may void this warranty

- Repair work undertaken without prior consent of Percival may void this warranty

- Damage caused by accident, fire, Force Majeure or Acts of God

Additional Terms

- This warranty is in lieu of any other warranties, expressed or implied, including merchantability or fitness for a particular purpose

- Warranty terms described herein cannot be changed without the expressed, written consent from Percival

- The owner agrees that Percival’s sole liability with respect to defective parts or materials shall be set forth in this warranty, and any claims for incidental or consequential damages are expressly excluded

- This warranty is non-transferable and applies only to the original purchaser

- Repaired or replaced parts shall be warranted only for the unexpired portion of the original warranty, unless otherwise defined

- Parts replaced under warranty shall become the property of Percival

To Make a Warranty Claim

- Document the issue with detailed notes and photographs along with the equipment’s serial number

- Contact Percival’s customer support team at customerservice@percival-scientific.com, parts@percival-scientific.com or 1-800-695-2743

- Once your claim is received, our customer support team will review it and contact you with further instructions

# Definitions

‘Controlled environment chamber’, ‘Incubator’, ‘Chamber’ and ‘Unit’ are all used interchangeably in this manual to describe the complete controlled environment chamber system manufactured by Percival Scientific, Inc.

# General

This manual is intended to provide instructions on the installation, operation, maintenance and troubleshooting of Percival Scientific’s I-Series biological incubator product line, including all I22, I36, I41 and I66 base models.

Biological incubators play a crucial role in various scientific and medical fields, serving as controlled environments for the growth and maintenance of biological samples. These versatile devices find extensive applications in microbiology, cell culture, and molecular biology research. In microbiology, incubators are used to cultivate and study microorganisms under optimal temperature, humidity, and atmospheric conditions. This is essential for identifying pathogens, testing antimicrobial agents, and developing new vaccines. In cell culture laboratories, incubators provide the ideal conditions for growing and maintaining cell lines, tissues, and organoids, supporting research in areas such as cancer biology, stem cell research, and drug discovery.

Beyond research applications, biological incubators are indispensable in clinical settings, biotechnology industries, and entomological studies. In hospitals and diagnostic laboratories, they are used for culturing patient samples to diagnose infections and determine antibiotic susceptibility. The pharmaceutical industry relies on incubators for drug development and quality control processes, including stability testing of new medications. In entomology, incubators are crucial for studying insect development, behavior, and physiology. They allow researchers to control environmental factors such as temperature, humidity, and light cycles, which are essential for rearing insects, investigating their life cycles, and conducting experiments on pest control methods. Entomologists use incubators to study insect responses to various stimuli, assess the effects of pesticides, and even explore the potential of insects in sustainable food production. Additionally, biological incubators find applications in food science for studying food safety and quality, as well as in environmental research investigating the impact of various conditions on microbial ecosystems. The versatility and precision of modern biological incubators continue to advance scientific discovery and progress across numerous disciplines, including the field of entomology.

## Equipment Overview:

(I22 Series)

(I36 Series)

(I41 Series)

(I66 Series)

## Product Specifications

### I22 Series Specification Table

| Model |  |  | I-22NL |  |  | I-22L |
| --- | --- | --- | --- | --- | --- | --- |
| Temperature Range | Lights Off | Low | 2.0°C |  |  |  |
|  |  | High | 44.0°C |  |  |  |
|  |  | Range | ±0.5°C |  |  |  |
|  | Lights On | Low | N/A | 4.0°C |  |  |
|  |  | High | 44.0°C |  |  |  |
|  |  | Range | ±0.5°C |  |  |  |
| Exterior Dimensions |  | Width | 33.5 in (85.1 cm) |  |  |  |
|  |  | Depth | 36.6 in (93.0 cm) |  |  |  |
|  |  | Height | 77.9 in (197.9 cm) |  |  |  |
| Interior Volume per Compartment |  | ft³ | 14.6 |  |  |  |
|  |  | m³ | 0.4 |  |  |  |
|  |  | L | 413.4 |  |  |  |
| Number of Tiers per Compartment |  |  | 3 |  |  | 2 |
| Shelves per Tier |  |  | 1 |  |  |  |
| Shelf Dimensions |  | Width | 28.8 in (73.0 cm) |  |  |  |
|  |  | Depth | 27.0 in (68.6 cm) |  |  |  |
| Total Shelf Area per Compartment |  | ft² (m²) | 16.2 (1.5) |  | 10.8 (1.0) |  |
| Growing Height per Tier |  | in (cm) | 11.0 (27.9) |  | 10.0 (25.4) |  |
| Door Opening |  | Width | 29.3 in (74.3 cm) |  |  |  |
|  |  | Height | 27.8 in (70.6 cm) |  |  |  |
| Light Intensity @ 6” (µmol/m²/s) |  |  | N/A |  | 120 |  |
| Electrical per Compartment |  | VAC | 120V |  |  |  |
|  |  | Phase | 1 |  |  |  |
|  |  | Frequency | 60Hz |  |  |  |
|  |  | Cords (NEMA) | 1 (5-15P) |  |  |  |
|  |  | RLA per cord | 8.9A |  | 9.3A |  |
|  |  | MCA | 11.1A |  | 11.6A |  |
| Heat Rejection to Ambient |  | Btu/h (W) | 1770 (519) |  | 1956 (573) |  |
| Shipping Weight |  | lb (kg) | 660 (300) |  |  |  |

### I36 Series Specification Table

| Model |  |  | I-36NL |  | I-36VL |  |  |  |  | I-36LL |  | I-36LLVL |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Temperature Range | Lights Off | Low | 2.0°C |  |  |  |  |  |  |  |  |  |
|  |  | High | 44.0°C |  |  |  |  |  |  |  |  |  |
|  |  | Range | ±0.5°C |  |  |  |  |  |  |  |  |  |
|  | Lights On | Low | N/A |  |  | 4.0°C |  | 5.0°C |  |  |  | 4.0°C |
|  |  | High | 44.0°C |  |  |  |  |  |  |  |  |  |
|  |  | Range | ±0.5°C |  |  |  |  |  |  |  |  |  |
| Exterior Dimensions |  | Width | 33.5 in (85.1 cm) |  |  |  |  |  |  |  |  |  |
|  |  | Depth | 33.6 in (85.4 cm) |  |  |  |  |  |  |  |  |  |
|  |  | Height | 77.2 in (196.1 cm) |  |  |  |  |  |  |  |  |  |
| Interior Volume |  | ft³ | 29.7 |  |  |  |  |  |  |  |  |  |
|  |  | m³ | 0.8 |  |  |  |  |  |  |  |  |  |
|  |  | L | 840.4 |  |  |  |  |  |  |  |  |  |
| Number of Tiers |  |  | 6 |  | 5 |  |  |  |  | 4 |  | 4 |
| Shelves per Tier |  |  | 1 |  |  |  |  |  |  |  |  |  |
| Shelf Dimensions |  | Width | 28.8 in (73.0 cm) |  |  |  |  |  |  |  |  |  |
|  |  | Depth | 27.0 in (68.6 cm) |  |  |  |  |  |  |  |  |  |
| Total Shelf Area |  | ft² (m²) | 32.3 (3.0) | 27.0 (2.5) |  |  |  |  | 21.6 (2.0) |  |  |  |
| Growing Height |  | in (cm) | 8.0 (20.3) | 9.5 (24.1) |  |  |  |  | 11.25 (28.6) |  | 11.25 (28.6) |  |
| Door Opening |  | Width | 29.3 in (74.3 cm) |  |  |  |  |  |  |  |  |  |
|  |  | Height | 57.5 in (146.1 cm) |  |  |  |  |  |  |  |  |  |
| Light Intensity @ 6” (µmol/m²/s) |  |  | N/A |  |  |  | 120 |  | 125 |  |  | 230 |
| Electrical |  | VAC | 120V |  |  |  |  |  |  |  |  |  |
|  |  | Phase | 1 |  |  |  |  |  |  |  |  |  |
|  |  | Frequency | 60Hz |  |  |  |  |  |  |  |  |  |
|  |  | Cords (NEMA) | 1 (5-15P) |  |  |  |  |  |  |  |  |  |
|  |  | RLA per cord | 9.0A | 9.5A |  |  |  |  | 9.9A |  | 10.5A |  |
|  |  | MCA | 11.3A | 11.9A |  |  |  |  | 12.4A |  | 13.2A |  |
| Heat Rejection to Ambient |  | Btu/h (W) | 1,865 (549) | 2,120 (624) |  |  |  |  | 2,224 (654) |  | 2,493 (733) |  |
| Shipping Weight |  | lb (kg) | 440 (200) | 448 (203) |  |  |  |  | 446 (202) |  | 452 (205) |  |

### I41 Series Specification Table

| Model |  |  | I-41NL |  | I-41VL |  |  |  |  | I-41LL |  | I-41LLVL |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Temperature Range | Lights Off | Low | 2.0°C |  |  |  |  |  |  |  |  |  |
|  |  | High | 44.0°C |  |  |  |  |  |  |  |  |  |
|  |  | Range | ±0.5°C |  |  |  |  |  |  |  |  |  |
|  | Lights On | Low | N/A |  |  | 4.0°C |  | 4.0°C |  |  |  | 5.0°C |
|  |  | High | 44.0°C |  |  |  |  |  |  |  |  |  |
|  |  | Range | ±0.5°C |  |  |  |  |  |  |  |  |  |
| Exterior Dimensions |  | Width | 41.0 in (104.1 cm) |  |  |  |  |  |  |  |  |  |
|  |  | Depth | 33.6 in (85.4 cm) |  |  |  |  |  |  |  |  |  |
|  |  | Height | 77.2 in (196.1 cm) |  |  |  |  |  |  |  |  |  |
| Interior Volume |  | ft³ | 37.7 |  |  |  |  |  |  |  |  |  |
|  |  | m³ | 1.06 |  |  |  |  |  |  |  |  |  |
|  |  | L | 1055.1 |  |  |  |  |  |  |  |  |  |
| Number of Tiers |  |  | 6 |  | 5 |  |  |  |  | 4 |  | 4 |
| Shelves per Tier |  |  | 1 |  |  |  |  |  |  |  |  |  |
| Shelf Dimensions |  | Width | 36.3 in (92.2 cm) |  |  |  |  |  |  |  |  |  |
|  |  | Depth | 27.0 in (68.6 cm) |  |  |  |  |  |  |  |  |  |
| Total Shelf Area |  | ft² (m²) | 40.8 (3.8) | 34.0 (3.2) |  |  |  |  | 27.2 (2.5) |  |  |  |
| Growing Height |  | in (cm) | 8.0 (20.3) | 9.5 (24.1) |  |  |  |  | 11.25 (28.6) |  | 11.25 (28.6) |  |
| Door Opening |  | Width | 36.8 in (93.5 cm) |  |  |  |  |  |  |  |  |  |
|  |  | Height | 57.5 in (146.1 cm) |  |  |  |  |  |  |  |  |  |
| Light Intensity @ 6” (µmol/m²/s) |  |  | N/A |  |  |  | 120 |  | 100 |  |  | 220 |
| Electrical |  | VAC | 120V |  |  |  |  |  |  |  |  |  |
|  |  | Phase | 1 |  |  |  |  |  |  |  |  |  |
|  |  | Frequency | 60Hz |  |  |  |  |  |  |  |  |  |
|  |  | Cords (NEMA) | 2 (5-15P) |  |  |  |  |  |  |  |  |  |
|  |  | Cord 1 RLA | 4.3A | 5.1A |  |  |  |  | 5.1A |  | 5.9A |  |
|  |  | Cord 2 RLA | 6.4A | 6.4A |  |  |  |  | 6.4A |  | 6.4A |  |
|  |  | MCA | 13.4A | 14.4A |  |  |  |  | 14.4A |  | 15.4A |  |
| Heat Rejection to Ambient |  | Btu/h (W) | 2,409 (709) | 2,662 (780) |  |  |  |  | 2,675 (784) |  | 2,952 (865) |  |
| Shipping Weight |  | lb (kg) | 523 (237) | 533 (242) |  |  |  |  | 533 (222) |  | 543 (246) |  |

### I66 Series Specification Table

| Model |  |  | I-66NL |  | I-66VL |  |  |  | I-66LL |  | I-66LLVL |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Temperature Range | Lights Off | Low | 2.0°C |  |  |  |  |  |  |  |  |
|  |  | High | 44.0°C |  |  |  |  |  |  |  |  |
|  |  | Range | ±0.5°C |  |  |  |  |  |  |  |  |
|  | Lights On | Low | N/A |  |  | 4.0°C |  |  |  |  |  |
|  |  | High |  |  |  | 44.0°C |  |  |  |  |  |
|  |  | Range |  |  |  | ±0.5°C |  |  |  |  |  |
| Exterior Dimensions |  | Width | 66.0 in (167.6 cm) |  |  |  |  |  |  |  |  |
|  |  | Depth | 33.6 in (85.4 cm) |  |  |  |  |  |  |  |  |
|  |  | Height | 77.2 in (196.1 cm) |  |  |  |  |  |  |  |  |
| Interior Volume |  | ft³ | 62.4 |  |  |  |  |  |  |  |  |
|  |  | m³ | 1.77 |  |  |  |  |  |  |  |  |
|  |  | L | 1768.1 |  |  |  |  |  |  |  |  |
| Number of Tiers |  |  | 6 |  | 5 |  |  |  | 4 |  | 4 |
| Shelves per Tier |  |  | 2 |  |  |  |  |  |  |  |  |
| Shelf Dimensions |  | Width | 28.8 in (73.2 cm) |  |  |  |  |  |  |  |  |
|  |  | Depth | 27.0 in (68.6 cm) |  |  |  |  |  |  |  |  |
| Total Shelf Area |  | ft² (m²) | 64.8 (6.0) | 54.0 (5.0) |  |  |  | 43.2 (4.0) |  |  |  |
| Growing Height |  | in (cm) | 8.0 (20.3) | 9.5 (24.1) |  |  |  | 11.25 (28.6) |  | 11.25 (28.6) |  |
| Door Opening (each) |  | Width | 29.1 in (73.9 cm) |  |  |  |  |  |  |  |  |
|  |  | Height | 57.5 in (146.1 cm) |  |  |  |  |  |  |  |  |
| Light Intensity @ 6” (µmol/m²/s) |  |  | N/A |  |  |  | 120 | 120 |  |  | 230 |
| Electrical |  | VAC | 120V |  |  |  |  |  |  |  |  |
|  |  | Phase | 1 |  |  |  |  |  |  |  |  |
|  |  | Frequency | 60Hz |  |  |  |  |  |  |  |  |
|  |  | Cords (NEMA) | 2 (5-15P) |  |  |  |  |  |  |  |  |
|  |  | Cord 1 RLA | 6.8A | 8.0A |  |  |  | 8.4A |  | 9.6A |  |
|  |  | Cord 2 RLA | 9.5A | 9.5A |  |  |  | 9.5A |  | 9.5A |  |
|  |  | MCA | 20.4A | 21.9A |  |  |  | 22.4A |  | 23.9A |  |
| Heat Rejection to Ambient |  | Btu/h (W) | 3206 (940) | 3583 (1050) |  |  |  | 3762 (1103) |  | 4341 (1272) |  |
| Shipping Weight |  | lb (kg) | 774 (351) | 790 (358) |  |  |  | 796 (361) |  | 815 (370) |  |

# Installation

To achieve and maintain optimum performance from the unit throughout its designed operating range(s), it is important to install the chamber according to the following guidelines.  Failure to follow the installation guidelines in this section may cause loss of environmental control, malfunction, premature system failure and may also void the manufacturer’s warranty.

## Site Preparation

The following considerations should be made when preparing a facility for installation of Percival controlled environments:

### Location

Locate the chamber in a clean area with a level surface and adequate space for door swings, ventilation, access of facility services and the unit’s service connections and access panels.

Prior to the unit’s delivery date, ensure there is a clear path from the delivery dock to final chamber location, taking doorways, elevators, corridor widths and turns into account.

Warning! 22, 36, 41 and 66 series chambers are top heavy.  Please use caution when removing a chamber from its shipping container and while moving it into position.  These tasks should be undertaken with care by more than one person.

Spatial requirements vary by chamber model and optional equipment.  Please refer to all supplementary installation instructions as optional equipment can change the space required for a chamber’s proper operation.

Generally, space should be planned for door swings, removal of service panels, condensing unit intake and exhaust ventilation and shutoff valves for chamber services.

Provide at least six inches of clearance on the side and top of the 36, 41, and 66 units to facilitate proper air flow through the condensing unit’s heat exchanger.  Air is drawn in through the right-handed side vent and exhausted through the top vent.

There are two chambers on a 22 unit, so there are two condensing units.  This design requires vents on the right and left for intake and one on the back side for exhaust. A minimum of six inches of clearance is required on the sides and behind the unit for proper air flow.

### Ambient Conditions

Published chamber performance specifications are based on the following ambient environmental conditions.  To ensure best performance of an I-Series controlled environment chamber, locate the unit in an environment where the following conditions can be maintained:

Temperature = 75 °F (23.8 °C)

Relative Humidity = 55 %RH

### Chamber Heat Rejection to Ambient

To achieve and maintain specified conditions within a controlled environment chamber, the unit’s cooling system must remove excess heat from within the chamber.  This excess heat is rejected from the system and must be accounted for when planning for a chamber’s installation.

The base I-Series chamber models’ estimated maximum heat rejection to ambient can be found in the Product Specifications section of this manual.

It is necessary to ensure the HVAC system where the chamber is installed has sufficient capacity to remove heat rejected by all chambers within the space.  Insufficient heat removal capacity will result in elevated ambient temperatures and possible degradation in chamber performance.

### Electrical Services

The I-Series controlled environment chamber is typically provided with one or more power cords that must be plugged into electrical building services of appropriate type, rating and ampacity.

Always verify a chamber’s electrical requirements against the provided electrical services prior to making final connections and energizing the unit.

A chamber’s electrical requirements are listed on its serial tag, which is located inside the unit’s door on the interior side wall of the controlled environment space.  These electrical requirements include the operating voltage, phase(s) and maximum running load amperage (RLA) for the chamber, per the following requirements:

- Supply voltage must be within ±10% of the serial tag value.

- The number of phases provided should match the serial tag value.

- The electrical service must have a minimum circuit ampacity (MCA) sufficient to support the chamber’s rated maximum running load amperage (RLA).  Service MCA should ultimately be determined by local authorities having jurisdiction (AHJ), according to state, local and building electrical codes.  As a rule of thumb, MCA can be estimated by the following equation:

MCA=1.25×RLAmax

- The provided electrical service should not be smaller than the MCA.  In many cases the service itself will be larger than the MCA. For example, a standard 30A electrical service might be specified to comply with a chamber’s MCA requirement of 27A.

If a chamber is provided with multiple power cords, each cord will have its own RLA/MCA requirement.  If both cords are to be connected on a single circuit with multiple receptacles, the service must be sized for the COMBINED RLA of both cords.  An example equation is provided below for reference:

MCAcombined=1.25 ×(RLAcord 1+ RLAcord 2)

In lieu of a single electrical service sized for multiple cords, connect each power cord to its own separate, dedicated electrical service.

Failure to follow the electrical service requirements outlined above can result in overloaded electrical services, tripped circuit protection on the supply and damage to facility and chamber electrical infrastructure such as wiring, plugs and receptacles.

Damage caused from incorrect or improperly sized electrical services is not covered under the manufacturer’s warranty.

The duplex receptacle shown above can physically accommodate two 20A rated cords, but may not have sufficient capacity to supply power at the full rating of two cords, simultaneously.  Verify the combined load the duplex receptacle’s electrical service can safely support before plugging in multiple power cords.

Do not plug multiple chamber power cords into a duplex receptacle unless the circuit is dedicated and verified to have an MCA sufficient to support the combined RLA of both connected power cords.

Check quadplex receptacles for proper labelling indicating the presence of multiple circuits prior to plugging in multiple chamber power cords

### Drainage and Plumbing Services

All base model I-Series controlled environments reject condensed water from their evaporator coils.  The amount of condensate is dependent upon a combination of controlled environment settings, ambient conditions, door opening frequency and duration, air exchange through (optional) fresh air ports, and characteristics of products placed inside the environment.

Optional relative humidity control systems increase the amount of condensate rejected from a given unit.

Whenever the evaporator coil operates at a temperature below the dew point inside the controlled environment, moisture present in the circulated air condenses onto the cooling coils.  This condensate sheds off the coils, is routed through a drain and expelled out of the environment’s floor drain.

Base model I-Series incubators are provided with one 7/8 in ID (1 in OD) flexible drain line.  Ideally, this drain line should be routed toward a floor drain close in proximity to the chamber to properly manage condensate rejected from the unit.

If the installation location for the chamber does not provide a nearby floor drain, the following drain accommodations can be considered:

- Extend the existing chamber drain tubing and route it to a floor drain that is farther away than the chamber’s footprint.

- Replace the existing chamber drain fitting and tubing with a Q19 (External Drip Pan) optional equipment package.

- Route existing chamber drain tubing to a Q19 (External Drip Pan) when no space is available directly below the chamber drain

- Connect the existing chamber drain system to a condensate pump

Please note that the standard drain system in I-Series controlled environment chambers works via gravity.  Condensate only drains properly if the facility drain is lower in elevation than the chamber floor drain and tubing.

## Unpackaging

### Inspect the Shipment

At the time of delivery, it is very important to verify that the proper quantity and type of products were received as per the Bill of Lading, and to inspect those products for shipping damage prior to accepting the shipment.

- Carefully remove all packaging materials that surround the unit, ensuring no loose chamber components are accidentally discarded.

- Visually inspect both the interior and exterior of the unit for dents, scratches or any other signs of damage

- It is extremely important to note any damage on the shipping receipt so that corrective actions can be taken as efficiently and expeditiously as possible.  The shipping documentation provides the following options:

  - Accept the shipment as-is

  - Accept the shipment with damages documented, understanding that corrective actions to rectify the damage will be performed onsite

  - Reject the shipment, where damage is sufficient to require the product to be returned to Percival, where corrective actions can be taken at the manufacturing facility

### Remove the Chamber from its Skid

Warning! All I-Series chambers are heavy.  22, 36, 41 and 66 series models are top-heavy.  Please use caution when removing a chamber from its shipping container and while moving it into position.  This task should be undertaken with care by more than one person.

- Remove shelving and any additional items packaged inside the chamber that can shift when the chamber is moved

- If there is a drain line affixed to the shipping skid, free it and temporarily secure it to the chamber to prevent it from hindering movement of unit

- Remove any metal shipping brackets that affix the chamber to its skid

- Remove any bolts that affix the chamber to its skid

- Remove the outer support boards on the sides of the chamber, allowing for more room to move the unit upon the skid

  - Remove wood screws

  - Slightly tip the unit to each side and pull the boards out from under the chamber

- Manipulate the chamber toward the back of the shipping skid

- Carefully allow the chamber to roll off the shipping skid backward, providing adequate support behind the unit to prevent it from tipping over.  The rear casters, located at the back of the unit should come off the skid first.

- Once the rear casters are in contact with the ground with the chamber in its tipped position, slide the skid out from under chamber.

- After removal of the skid, carefully allow the chamber to tip forward, allowing all four casters to rest on the floor.

### Moving the Chamber

Warning! All I-Series chambers are heavy.  I22, I36, I41 and I66 series models are top heavy.  Please use caution while moving any chamber.  This task should be undertaken with care by more than one person.

Do not attempt to move any controlled environment chamber loaded with experiments, samples or other sensitive materials.

I22, I36, I41 and I66 Series incubators are provided with swivel casters for the specific purposes of rolling empty units into their installation location, and moving units as needed for service and maintenance efforts.

The caster system is designed for moving empty chambers over short distances.  Special care is needed when moving a unit over uneven floor surfaces.

## Installation Guidelines

Once placed in a suitable location as prescribed in the preceding sections of this user manual, proceed with the following installation steps:

### Level the Unit

Locate I-Series incubators upon a level surface to facilitate proper function and drainage.

Adjustable levelers are provided with each base model I22, I36, I41 and I66 Series chamber as a means of levelling the chamber for experiments and preventing chambers from rolling unexpectedly when in use.

Use a crescent wrench to adjust the vertical position of each leveler as necessary.  A clockwise turn extends the leveler toward the floor and a counterclockwise turn retracts the leveler.

Levelers are designed to support the entire weight of a chamber, so can be adjusted to lift the unit off its casters to prevent inadvertent movement, if desired.

### Inspect the Control Box

While unplugged from electrical services, open the chamber’s control panel and perform a visual inspection of the electrical components installed within it.

- Check the state of the resettable circuit breakers and reset any found in an ‘open’ state

- Check for any components that are not physically secured to the components mounting plate, control box housing or top of the insulated box.

- Check for any connector plugs that are dislodged or not fully seated into their receptacles and re-secure them as needed with a firm press.

- Check for any obviously loose or disconnected wiring

If any components have become dislodged or there is any question as to where a disconnected plug or wire should be connected, do not proceed any further with chamber installation or operation and contact Percival’s customer support team at 1-800-695-2743 or customerservice@percival-scientific.com

### Install Shelving and Lights

All I-Series controlled environment chambers are provided with steel wire shelving, which is shipped uninstalled and packaged within each chamber.  To install shelving in an I-Series base model chamber:

- Unpackage each shelf and locate shelf hardware (pilaster clips)

- Install pilaster clips:

  - Four clips, located at the same height are required to properly support each shelf

  - Numbers are embossed into each pilaster to assist with alignment of the clips

  - Insert the upper pilaster clip tab into the pilaster first

  - Twist the pilaster clip in a downward motion and insert the lower tab into the lower pilaster slot.

Note:  It may be necessary to pinch the clip slightly, upon insertion, to obtain proper engagement of both clip tabs into the pilaster.

The image below shows an example of a properly installed pilaster clip:

- Place shelves onto the installed pilaster clips, ensuring that the vertical tab on each clip sits inboard of the shelf frame, securing each shelf in place.  When properly installed a shelf can only be removed by lifting it vertically off its clips.

The image below shows an example of a properly installed wire shelf:

For base model I-Series incubators provided with IncuWhite™ (LL, VL and LLVL), please refer to the supplemental IW User Manual for instruction on the installation of horizontal lighting.

Vertical IncuWhite™ lighting is shipped pre-installed in VL and LLVL base model incubators.

### Connect Services

- Plug the unit into properly sized electrical services, ensuring that all plugs are completely and securely inserted

- Route the chamber’s drain to a floor drain or other drainage accommodation

- Connect services for optional equipment as directed in supplemental operation manuals (e.g. demineralized water, compressed gases, chilled water, ventilation ductwork, etc...)

## Startup and Commissioning

### Turn on the Chamber

Set the main power switch located on the left side of the control panel to the ‘On’ or ‘|’ position to energize the chamber.  The IntellusUltra controller takes approximately 15-20 seconds to start up.  Once the controller has started up, and there are no active alarm conditions, the rest of the chamber components will energize.

### IntellusUltra Control System

The primary method of operating an I-Series controlled environment chamber is through the C8 IntellusUltra control system’s keypad overlay. The IntellusUltra overlay (shown below) is located on the exterior of the I-Series base model’s control panel.

Refer to the IntellusUltra Quick Guide provided with the chamber for basic instructions on the controller’s operation.  This guide can also be downloaded at https://www.percival-scientific.com/wp-content/uploads/2023/03/IntellusUltra-Quick-Guide.pdf

Refer to the C8 Control System User Manual for comprehensive operational instructions for the chamber’s control system

If the primary control system user interface does not look like the image above, then the chamber is equipped with an optional touch screen user interface.

Please refer to the supplementary Touch Screen User Manual for instructions on the operation of this optional system.

When first energized the IntellusUltra control system is set to operate in Manual Mode, and has the following factory settings:

| Control Feature | Factory Setting |
| --- | --- |
| System Time | Current North American Central Time Zone |
| Operation Mode | Manual |
| Programming Step Type | Non-ramping |
| Temperature Set Point | 25°C |
| Temperature Limit 1 High Alarm | 35°C |
| Temperature Limit 1 Low Alarm | -5°C |
| Temperature Limit 2 High Alarm | 38°C |
| Temperature Limit 2 Low Alarm | -8°C |
| Light Output 1 (LL, VL and LLVL models) | 0% (Off) |
| Light Output 2 (LLVL models) | 0% (Off) |

In general, any optional features controlled via the IntellusUltra are disabled by default.  Please refer to the user manuals provided with optional equipment for information pertaining to their factory settings, setup and operation.

### Calibration

All controlled environment chambers from Percival Scientific must pass rigorous quality checks and validation testing prior to shipment.  As part of the validation process, all chamber control systems are calibrated against an NIST standard device at the following conditions:

| Condition | Calibration |
| --- | --- |
| Temperature w/ Lights Off | 25°C |
| Temperature w/ Lights On (for lighted models) | 25°C |
| Relative Humidity w/ Lights Off (if RH option is equipped) | 75%RH @ 25°C |
| Relative Humidity w/ Lights On (if RH option is equipped) | 75%RH @ 25°C |

At these calibration points, any discrepancy between an IntellusUltra sensor (located by the intake of the evaporator coil) reading and the NIST calibration standard (located in the center of the chamber environment) are entered into the control system as calibration offsets.

For any given environmental parameter controlled by the chamber:

Calibration Offset=Calibration Standard Value-Control System Process Value

For example, if the control system reads a temperature process value of 25°C and the calibration standard reads 24.5°C, then an offset of -0.5°C is saved into the controller.  For each condition, the measured offset is applied to the process value so that it matches the calibration standard.

TOffset= TStandard- TProcess Value

TOffset= 24.5°C- 25.0°C

TOffset= -0.5°C

Any offsets with lights turned off are called Night Offsets and any offsets with any amount of lights turned on are called Day Offsets.

Please refer to the provided control system user manual for additional information regarding calibration offsets.

Factory calibration offsets serve as suitable values for general chamber use.  Consider recalibrating the chamber if precise environmental control is required at specific conditions that are significantly different than the factory calibration points.

# Operation

## Guidelines for Loading a Chamber

When placing products, specimens or experiments onto shelves inside an I-Series controlled environment, adhere to the following guidelines to ensure the unit operates at peak performance:

- For best temperature uniformity, evenly distribute items throughout the chamber space and allow space between items to promote proper air flow

- As a good rule of thumb, do not populate more than 80% of any shelf’s surface area

- Do not place more than 85 lbs (38.6 kg) onto any I41 shelf

- Do not place more than 125 lbs (56.8 kg) onto any I22, I36 or I66 shelf

- Do not place more than 500 lbs (227.3 kg) total in a chamber

Avoid placing items in a unit that produce significant amounts of heat.  I-Series base model chambers are designed to accommodate a product heat load of 2 W/ft² of shelf space.  Any heat introduced into the chamber environment beyond this amount can result in loss of environmental control.

## Running a Control Profile

- Close the main chamber door, ensuring a proper seal between the door gasket and insulated box

- Configure static or multi-step profile process value set points

- Check alarm limits, ensuring that they are set above the maximum and below the minimum process values.  Limits should not be set closer than 2-3°C above or below the operational range of the entire control profile.

- Select the mode of operation on the control system (e.g. Manual, Diurnal, Program or Sequence).

Once an operating mode is set, the IntellusUltra control system automatically executes any set points specified in the profile.

If an alarm threshold is exceeded, the controller produces an audible alarm until the alarm is acknowledged by the user.  The alarm state remains active until the source of the alarm is resolved.  A temperature alarm condition shuts down all chamber sub-systems inside the environment until the alarm state is cleared.  Whenever process values fall back within allowable limits, the control system automatically reverts to normal operation, requiring no involvement from the user.

The IntellusUltra controller has an internal battery which allows it to keep time in the event of a power cycle or outage.  Upon power up, the chamber will resume the previously executed control profile.

- Control is resumed around static process value set points when the controller is in manual mode

- When executing a multi-step program, the controller resumes control around the current program step as determined by the controller system time.

- When executing a sequence of programs, the controller resumes control around the current program iteration and step as determined by the system time unless power is off at 12:00 am.  If power is off at 12:00am the program iteration count will not be incremented.

For best control, minimize the frequency and duration of chamber door openings.

# Maintenance

Performing regular maintenance on an environmental chamber reduces performance degradation, improves reliability and extends operational life.  In addition to any user-specific maintenance protocols, the following general maintenance tasks should be performed at the recommended intervals.

Turn chamber power off prior to performing any maintenance inside the environment.

Unplug or disconnect chamber supply power prior to performing any maintenance inside the chamber’s mechanical compartment or control box.

## General Maintenance Summary Table

| Maintenance Task | Recommended Interval Schedule |
| --- | --- |
| Clean the Chamber | Every 3 months, or as needed |
| Clean the Condenser | Every 3 months, or as needed |
| Clean the Evaporator | Every 12 months |

## Cleaning the Chamber

Wipe down all interior and exterior surfaces with a solution of water and mild detergent.  Do not use abrasive agents or cleaning pads as these could scratch surfaces.

Glass cleaners can be used on stainless steel surfaces to prevent streaking.

## Clean the Heat Exchangers

Air must pass freely through the evaporator and condenser for proper performance.  Ambient air is drawn into and forced through the condenser by the condensing unit fan.  Chamber process air is drawn into and forced through the evaporator by a bank of fans inside the evaporator cover.  In both cases, the fans run continuously while the chamber is in operation and there are no active alarms.

Power down the chamber prior to any cleaning effort on refrigeration system components.

Please be aware that some parts of the condensing unit become hot while the unit is in operation and can remain hot for some time after the unit is powered down.

### Clean the Condenser

Once powered down, remove any dust or debris collected upon the condenser fins, copper tubing, fan blades and motor.  Blow dust out of the condenser from behind with compressed air, a brush, or vacuum and wipe dust from the fan blades.

To access the condenser coil, remove the vented cover panel on the right side of the unit’s mechanical section.

To access the condenser fan in an I22, I36, I41 or I66 Series incubator, remove the mechanical section’s top cover(s) and/or rear panel(s).

### Clean the Evaporator

Once the unit is powered down, remove the evaporator cover to access the evaporator coil and air circulation fans.

- Disconnect the flexible condensate drain line from the evaporator housing

- Remove the sheet metal screws that secure the evaporator cover to the main housing or support brackets, depending on the incubator model.

- Disconnect the electrical connection(s) between the fans affixed to the cover and main housing

A picture of the exposed top-mounted evaporator coil is shown below for reference. This style of coil is utilized in 22, 36, 41 and 66 Series incubators.

(The evaporator above is shown with optional item: Q9 (Phenolic Coating))

Clean the evaporator coil by dislodging dirt and debris via soft brush, compressed air or vacuum, and inspect the condensate drain line for any cracks, wear or blockages.

Remove dust and debris from the evaporator fans and wipe any dirt or residue from inside the evaporator cover and housing.

If the evaporator has a brown phenolic coating and/or a white painted housing as shown in the photo above, please refer to the supplemental Q9 Phenolic Coating User Manual for specific cleaning requirements in lieu of these general guidelines.

# Troubleshooting

## Loss of Temperature Control – General

| Possible Reason | Solution |
| --- | --- |
| Improper or incorrect signal between controller and control relay(s) | Check for ~5VDC output signal between the controller and solid-state relay. Output state is indicated by LEDs on the controller overlay. Contact factory if insufficient output voltage from the control system |
| Temperature sensor | Check the sensor for correct output signal. Refer to the RTD temperature vs. resistance table in the controller manual. Contact factory if sensor output does not match the table values. |
| Solid-state relays | If correct trigger voltage is correct to the relay coil, check for the correct output voltage through the relay. Refer to the chamber’s electrical schematic for the VAC output to the affected heating and/or cooling device(s) |
| Hot gas solenoid valve | When in heating mode, the hot gas solenoid should be in the open position allowing hot gas to flow into the evaporator coil. Check the solenoid valve for proper actuation and the valve body for internal blockages. |
| Liquid line solenoid valve | When in cooling mode, the liquid line solenoid valve should be in the open position, allowing liquid refrigerant to flow into the thermal expansion device(s). Check the solenoid for proper actuation and the valve body for internal blockages. |

## Loss of Temperature Control - Overheating

| Loss of refrigerant | Check system pressures and check for leaks with a leak detector. Evacuate the system, repair leak(s) and recharge the system. Refer to the unit’s refrigeration diagram and serial label. |
| --- | --- |
| Restriction in the refrigeration piping or componentry | If adequate pressure and no leaks, there may be a restriction in the unit’s liquid line solenoid valve, filter/drier or thermal expansion device. Check the line temperature directly before and after each device. There may be a blockage if the outlet is colder than the inlet to the solenoid and drier. |
| Insufficient heat exchange through the condenser | Check the ambient air temperature is below 90°F. Make sure the condenser fan and coils are clean. Make sure the condenser fan is running. |
| Compressor motor fails to pump refrigerant throughout the system | Check electrical power to the motor and start components. Perform an ohm check on the motor windings. Check high-side system pressure if the unit runs |
| Insufficient heat exchange through the evaporator | Check for proper fan operation. Fans should be continuously running, drawing air into the evaporator. Make sure the evaporator coils are clean. |

## Loss of Temperature Control – Overcooling

| Electric heater malfunction | Check output voltage from the controller to the heater relay. Check output voltage from the heater relay to the heater element. Check the circuit breaker in the heater electrical circuit. |
| --- | --- |
| Malfunctioning hog gas solid-state relay (capillary tube systems) | Check output voltage from the controller to the hot gas line solenoid valve. Check voltage between the relay and solenoid valve. The solenoid valve should be open whenever the ‘heat’ LED is lit |
| Malfunctioning liquid solid-state relay (TXV systems) | Check output voltage between controller and relay. Check voltage between relay and solenoid. The solenoid should fully close when the ‘heat’ LED indicator light is energized. |
| Blocked hot gas solenoid valve (capillary tube systems) | If the hot gas solenoid valve gets blocked, then flow will divert through the capillary tube, causing the system to cool even when calling for heat. |

## General

| Chamber does not power on | Check facility electrical services and circuit breakers. Check that the chamber is properly plugged into the facility services. Check circuit breakers inside the unit’s control box. |
| --- | --- |
| Individual chamber component does not turn on | Check controller settings and verify that the device has been properly enabled. Check the control box for tripped circuit breakers. |
| Condensation | The amount of condensation that can form inside and be rejected from a controlled environment can fluctuate, and is highly dependent upon ambient conditions, interior environmental conditions, fresh air exchange through (optional) fresh air ports, frequency of door openings, product load and optional humidity control systems. Minimize door openings and the amount of moisture present inside the unit, and close the access port |

## Parts and Service

Replacement parts can be ordered from Percival Scientific at:

Ph: 1-800-695-2743

Email: parts@percival-scientific.com

For additional questions regarding this product, please contact Percival’s customer support team at:

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com