PGC-15 Chamber

Operator’s Manual

505 Research Drive  Perry, Iowa 50220

Toll-free 800.695.2743  515.465.9363  fax 515.465.9464

www.percival-scientific.com

[This page intentionally left blank]

# Introduction

To help customers create better science, Percival Scientific continuously studies and refines the design of its products to provide the best possible solutions and equipment in the industry.  As such, changes in the design and/or operation of equipment may occur without publication.  Please feel free to contact Percival Scientific directly for product updates.

Please read this manual carefully before proceeding with operation

# Revision Log

| Version | Date | Description |
| --- | --- | --- |
| Rev A | 02/05/2025 | Initial Release |

Introduction	3

Revision Log	4

Warranty Information	7

Definitions	9

General	9

Equipment Overview	10

Product Specifications	11

Installation	12

Site Preparation	12

Location	12

Ambient Conditions	13

Chamber Heat Rejection to Ambient	13

Electrical Services	13

Drainage and Plumbing Services	15

Unpackaging	16

Inspect the Shipment	16

Remove the Chamber from its Skid	17

Moving the Chamber	17

Installation Guidelines	18

Level the Unit	18

Inspect the Control Box	18

Release Light Fixture Counterweight	19

Connect Services	19

Startup and Commissioning	19

Turn on the Chamber	19

IntellusUltra Control System	20

Calibration	21

Operation	22

Guidelines for Loading a Chamber	22

Fresh Air Exchange	23

Running a Control Profile	24

Maintenance	25

General Maintenance Summary Table	25

Cleaning the Chamber	25

Clean the Bench Flooring & Air Diverters	26

Clean the Evaporator Coil(s)	27

Clean the Floor & Floor Drain	28

Clean the Heat Exchangers	28

Clean the Condenser	29

Clean the Aspirator Box Fan	30

Troubleshooting	30

Loss of Temperature Control – General	30

Loss of Temperature Control - Overheating	31

Loss of Temperature Control – Overcooling	31

General	31

Parts and Service	32

# Warranty Information

Equipment Limited Warranty

Thank you for purchasing equipment from Percival Scientific, Inc.  Our goal is to provide every customer with an exceptional ownership experience.  We are committed to providing products and services of the highest quality, and your satisfaction is our top priority.  In support of this objective, we present the following warranty.

Warranty Terms

Percival Scientific, Inc. (hereinafter referred to as Percival) warrants that equipment manufactured by Percival shall be free from defect in material or workmanship, when installed, operated and maintained in accordance with instructions provided by Percival.

Warranty Coverage

- Coverage period shall be twenty-four (24) months from the date of delivery

- Percival’s obligation under this warranty is limited to the repair or replacement of defect at no charge for parts, shipping and labor

- Percival reserves the right to inspect the equipment after a warranty claim is made to determine the proper course of corrective action

Exclusions

- Parts or elements damaged from improper installation, including mechanical, electrical and plumbing services, IT infrastructure, inadequate room ventilation, improper water quality, and ambient conditions that are out of specification

- Parts or elements damaged from improper operation

- Parts or elements damaged from improper maintenance

- Parts or elements considered to be consumable, that outlive their normal expected terms of usefulness, such as: light bulbs, filters, desiccant material and CO2 scrubbing media)

- Damage resulting from shipment to or relocation within the customer’s facility after equipment has been signed for and ownership is transferred to the customer

- Modification to equipment by the customer without Percival’s consent may void this warranty

- Repair work undertaken without prior consent of Percival may void this warranty

- Damage caused by accident, fire, Force Majeure or Acts of God

Additional Terms

- This warranty is in lieu of any other warranties, expressed or implied, including merchantability or fitness for a particular purpose

- Warranty terms described herein cannot be changed without the expressed, written consent from Percival

- The owner agrees that Percival’s sole liability with respect to defective parts or materials shall be set forth in this warranty, and any claims for incidental or consequential damages are expressly excluded

- This warranty is non-transferable and applies only to the original purchaser

- Repaired or replaced parts shall be warranted only for the unexpired portion of the original warranty, unless otherwise defined

- Parts replaced under warranty shall become the property of Percival

To Make a Warranty Claim

- Document the issue with detailed notes and photographs along with the equipment’s serial number

- Contact Percival’s customer support team at customerservice@percival-scientific.com, parts@percival-scientific.com or 1-800-695-2743

- Once your claim is received, our customer support team will review it and contact you with further instructions

# Definitions

‘Controlled environment chamber’, ‘Plant growth chamber’, ‘Chamber’ and ‘Unit’ are all used interchangeably in this manual to describe the complete controlled environment chamber system manufactured by Percival Scientific, Inc.

# General

This manual is intended to provide instructions on the installation, operation, maintenance and troubleshooting of Percival Scientific’s PGC-15 controlled environment chamber.

Plant growth chambers play a crucial role in various scientific and medical fields, serving as controlled environments for studies in plant growth and development.  Their ability to tightly control and reproduce environmental factors including temperature, light, humidity, and carbon dioxide make them excellent for applications of plant research, plant pathology, seed germination, and crop production studies.  The PGC-15 is designed to grow plant species requiring high light intensity to maturity (i.e. more than 1,000 µmol/m²/s).  The work area and growing height of this model facilitates the growth of some larger plant varieties to maturity.  Cereals, citrus, grasses, soybeans, tomatoes, cotton, peanuts, and potatoes are commonly studied with Percival’s PGC Series chambers.

## Equipment Overview

## Product Specifications

| Model |  |  | PGC-15 |  |
| --- | --- | --- | --- | --- |
| Temperature Range | Lights Off | Low | 2.0°C |  |
|  |  | High | 44.0°C |  |
|  |  | Range | ±0.5°C |  |
|  | Lights On | Low | 7.0°C |  |
|  |  | High | 44.0°C |  |
|  |  | Range | ±0.5°C |  |
| Exterior Dimensions |  | Width | 95.3 in (242.2 cm) |  |
|  |  | Depth | 38.8 in (98.6 cm) |  |
|  |  | Height | 78.8 in (200.0 cm) |  |
| Interior Volume |  | ft³ | 94.8 |  |
|  |  | m³ | 2.7 |  |
|  |  | L | 2,683.4 |  |
| Number of Tiers |  |  | 1 |  |
| Shelves per Tier |  |  | N/A |  |
| Work Area Dimensions |  | Width | 73.0 in (185.4 cm) |  |
|  |  | Depth | 29.8 in (75.6 cm) |  |
| Total Shelf Area |  | ft² (m²) | 15.1 (1.4) |  |
| Growing Height |  | in (cm) | 48.0 (121.9) |  |
| Door Opening |  | Width | 25.8 in (65.6 cm) |  |
|  |  | Height | 48.3 in (122.7 cm) |  |
| Light Intensity @ 6” (µmol/m²/s) |  |  | 1,520 |  |
| Electrical |  | VAC | 120/230 |  |
|  |  | Phase | 1 |  |
|  |  | Frequency | 60Hz |  |
|  |  | Cords (NEMA) | N/A |  |
|  |  | RLA per leg | 13.2A | 12.3A |
|  |  | MCA | 16.5A |  |
| Heat Rejection to Ambient |  | Btu/h (W) | 5,187 (1,526) |  |
| Shipping Weight |  | lb (kg) | 1,350 (612) |  |

# Installation

To achieve and maintain optimum performance from the unit throughout its designed operating range(s), it is important to install the chamber according to the following guidelines.  Failure to follow the installation guidelines in this section may cause loss of environmental control, malfunction, premature system failure and may also void the manufacturer’s warranty.

## Site Preparation

The following considerations should be made when preparing a facility for installation of Percival controlled environments:

### Location

Locate the chamber in a clean area with a level surface and adequate space for door swings, ventilation, access of facility services and the unit’s service connections and access panels.

Prior to the unit’s delivery date, ensure there is a clear path from the delivery dock to final chamber location, taking doorways, elevators, corridor widths and turns into account.

Warning! The PGC-15 chamber is heavy.  Please use caution when removing the unit from its shipping container and while moving it into position.  These tasks should be undertaken with care by more than one person.

Spatial requirements vary by chamber model and optional equipment.  Please refer to all supplementary installation instructions as optional equipment can change the space required for a chamber’s proper operation.

Space required for installation of the standard PGC-15 controlled environment chamber is shown below:

### Ambient Conditions

Published chamber performance specifications are based on the following ambient environmental conditions.  To ensure best performance of the PGC-15 controlled environment chamber, locate the unit in an environment where the following conditions can be maintained:

Temperature = 75 °F (23.8 °C)

Relative Humidity = 55 %RH

### Chamber Heat Rejection to Ambient

To achieve and maintain specified conditions within a controlled environment chamber, the unit’s cooling system must remove excess heat from within the chamber.  This excess heat is rejected from the system and must be accounted for when planning for a chamber’s installation.

The PGC-15 chamber model’s estimated maximum heat rejection to ambient can be found in the Product Specifications section of this manual.

It is necessary to ensure the HVAC system where the chamber is installed has sufficient capacity to remove heat rejected by all chambers within the space.  Insufficient heat removal capacity will result in elevated ambient temperatures and possible degradation in chamber performance.

### Electrical Services

The PGC-15 controlled environment chamber is typically provided with a field wiring terminal block that must be direct wired into electrical building services of appropriate type, rating and ampacity.

Always verify a chamber’s electrical requirements against the provided electrical services prior to making final connections and energizing the unit.

A chamber’s electrical requirements are listed on its serial tag, which is located inside the unit’s door on the interior side wall of the controlled environment space.  These electrical requirements include the operating voltage, phase(s) and maximum running load amperage (RLA) for the chamber, per the following requirements:

- Supply voltage must be within ±10% of the serial tag value.

- The number of phases provided should match the serial tag value.

- The electrical service must have a minimum circuit ampacity (MCA) sufficient to support the chamber’s rated maximum running load amperage (RLA).  Service MCA should ultimately be determined by local authorities having jurisdiction (AHJ), according to state, local and building electrical codes.  As a rule of thumb, MCA can be estimated by the following equation:

MCA=1.25×RLAmax

- The provided electrical service should not be smaller than the MCA.  In many cases the service itself will be larger than the MCA. For example, a standard 30A electrical service might be specified to comply with a chamber’s MCA requirement of 27A.

If a chamber is provided with multiple power cords, each cord will have its own RLA/MCA requirement.  If both cords are to be connected on a single circuit with multiple receptacles, the service must be sized for the COMBINED RLA of both cords.  An example equation is provided below for reference:

MCAcombined=1.25 ×(RLAcord 1+ RLAcord 2)

In lieu of a single electrical service sized for multiple cords, connect each power cord to its own separate, dedicated electrical service.

Failure to follow the electrical service requirements outlined above can result in overloaded electrical services, tripped circuit protection on the supply and damage to facility and chamber electrical infrastructure such as wiring, plugs and receptacles.

Damage caused from incorrect or improperly sized electrical services is not covered under the manufacturer’s warranty.

The duplex receptacle shown above can physically accommodate two 20A rated cords.  Per the above RLA/MCA calculation, the maximum electrical load that can be applied to a 20A circuit is 16A. If the provided electrical service is rated for 16 amps total, the duplex cannot support two power cords drawing a combined load of greater than 16A.

Do not plug multiple chamber power cords into a duplex receptacle unless the circuit is dedicated and verified to have an MCA sufficient to support the combined RLA of both connected power cords.

Check quadplex receptacles for proper labelling indicating the presence of multiple circuits prior to plugging in multiple chamber power cords

### Drainage and Plumbing Services

The PGC-15 controlled environment rejects condensed water from its evaporator coil.  The amount of condensate is dependent upon a combination of controlled environment settings, ambient conditions, door opening frequency and duration, and characteristics of products placed inside the environment.

Optional relative humidity control systems increase the amount of condensate rejected from a given unit.

Whenever the evaporator coil operates at a temperature below the dew point inside the controlled environment, moisture present in the circulated air condenses onto the cooling coils.  This condensate sheds off the coil, is collected on the environment’s floor, then naturally expels out of the environment’s floor drain.

PGC-15 controlled environment chambers are provided with a 30 in (76.2 cm) length, 7/8 in ID (1 in OD) flexible drain line.  Ideally, this drain line should be routed toward a floor drain close in proximity to the chamber to properly manage condensate rejected from the unit.

If the installation location for the chamber does not provide a nearby floor drain, the following drain accommodations can be considered:

- Extend the existing chamber drain tubing and route it to a floor drain that is farther away than the chamber’s footprint.

- Replace the existing chamber drain fitting and tubing with a Q19 (External Drip Pan) optional equipment package.

- Connect the existing chamber drain system to a condensate pump

Please note that the standard drain system in a PGC-15 controlled environment chamber works via gravity.  Condensate will only drain properly if the facility drain is lower in elevation than the chamber floor drain and tubing.

## Unpackaging

### Inspect the Shipment

At the time of delivery, it is very important to verify that the proper quantity and type of products were received as per the Bill of Lading, and to inspect those products for shipping damage prior to accepting the shipment.

- Carefully remove all packaging materials that surround the unit, ensuring no loose chamber components are accidentally discarded.

- Visually inspect both the interior and exterior of the unit for dents, scratches or any other signs of damage

- It is extremely important to note any damage on the shipping receipt so that corrective actions can be taken as efficiently and expeditiously as possible.  The shipping documentation provides the following options:

  - Accept the shipment as-is

  - Accept the shipment with damages documented, understanding that corrective actions to rectify the damage will be performed onsite

  - Reject the shipment, where damage is sufficient to require the product to be returned to Percival, where corrective actions can be taken at the manufacturing facility

### Remove the Chamber from its Skid

Warning! The PGC-15 chamber is heavy.  Please use caution when removing the unit from its shipping container and while moving it into position.  These tasks should be undertaken with care by more than one person.

A forklift may be used to remove the PGC-105 from its skid, but the following considerations should be taken:

- The chamber’s weight is biased toward the mechanical control section

- The chamber’s mechanical end section must be supported when the unit is lifted.  The mechanical section’s fasteners are used to affix it to the main chamber when in rest upon a level surface but are not suitable to support its weight in shear.

- Do not place forks directly under the chamber’s floor drain

In the event a forklift is unavailable, the chamber can be removed from its skid by hand.

- Remove the metal shipping brackets that affix the chamber to the skid

- Remove any bolts that affix the chamber to the skid

- Manipulate the chamber toward the back of the shipping skid

- Carefully allow the chamber to roll off the shipping skid backward, providing adequate support behind the unit to prevent it from tipping over.  The rear casters, located at the back of the unit should come off the skid first.

- Once the rear casters are in contact with the ground with the chamber in its tipped position, slide the skid out from under chamber.

- After removal of the skid, carefully allow the chamber to tip forward, allowing all eight casters to rest on the floor.

### Moving the Chamber

Swivel casters are provided with PGC-15 controlled environment for the specific purposes of rolling empty units into their installation location, and moving units as needed for service and maintenance efforts.

The caster system is designed for moving empty chambers over short distances.  Do not attempt to move any PGC-15 chamber loaded with experiments, samples or other sensitive materials.

Warning! The PGC-15 chamber is heavy.  Please use caution when removing the unit from its shipping container and while moving it into position.  These tasks should be undertaken with care by more than one person.

## Installation Guidelines

Once placed in a suitable location as prescribed in the preceding sections of this user manual proceed with the following installation steps:

### Level the Unit

Adjustable levelers are provided with each PGC-15 chamber and allow for compensation of unevenness in the floor area and as a means of preventing chambers from rolling unexpectedly when in use.

Use a crescent wrench to adjust the vertical position of each leveler as necessary.  A clockwise turn extends the leveler toward the floor and a counterclockwise turn retracts the leveler.

Levelers are designed to support the entire weight of a chamber, so can be adjusted to lift the unit off its casters to prevent inadvertent movement, if desired.

### Inspect the Control Box

While unplugged from electrical services, open the chamber’s control panel and perform a visual inspection of the electrical components installed within it.

- Check the state of the resettable circuit breakers and reset any found in an ‘open’ state

- Check for any components that are not physically secured to the components mounting plate, control box housing or top of the insulated box.

- Check for any connector plugs that are dislodged or not fully seated into their receptacles and re-secure them as needed with a firm press.

- Check for any obviously loose or disconnected wiring

If any components have become dislodged or there is any question as to where a disconnected plug or wire should be connected, do not proceed any further with chamber installation or operation and contact Percival’s customer support team at 1-800-695-2743 or customerservice@percival-scientific.com

### Release Light Fixture Counterweight

PGC-15 controlled environment chambers are provided with an adjustable light fixture assembly controlled by a counterweight installed behind a wall cover panel.  The counterweight is secured with screws and washers to prevent movement during shipment.  Remove these screws and washers once the chamber has been moved into its position of use.

Remove screws & washers

### Connect Services

- Direct wire the unit into properly sized electrical services, ensuring that all field conductors are completely and securely inserted into the control box’s field wiring terminal block

- Route the chamber’s drain to a floor drain or other drainage accommodation

- Connect services for optional equipment as directed in supplemental operation manuals (e.g. demineralized water, compressed gases, chilled water, ventilation ductwork, etc...)

## Startup and Commissioning

### Turn on the Chamber

Set the main power switch located on the front of the control panel to the ‘On’ or ‘|’ position to energize the chamber.  The IntellusUltra controller takes approximately 15-20 seconds to start up.  Once the controller has started up, and there are no active alarm conditions, the rest of the chamber components will energize.

### IntellusUltra Control System

The primary method of operating a PGC-15 controlled environment chamber is through the C8T IntellusUltra control system’s touch screen user interface. The IntellusUltra touch screen (shown below) is located on the exterior of the PGC-15’s control panel.

Please refer to the supplementary Touch Screen User Manual for instructions on the operation of this user interface.

Additionally, the PGC-15 controlled environment chamber can be operated through the C8 IntellusUltra control system’s keypad overlay.  The IntellusUltra overlay (shown below) is located on the interior of the PGC-15’s control panel.

Refer to the IntellusUltra Quick Guide provided with the chamber for basic instructions on the controller’s operation.  This guide can also be downloaded at https://www.percival-scientific.com/wp-content/uploads/2023/03/IntellusUltra-Quick-Guide.pdf

Refer to the C8 Control System User Manual for comprehensive operational instructions for the chamber’s control system

When first energized the IntellusUltra control system is set to operate in Manual Mode, and has the following factory settings:

| Control Feature | Factory Setting |
| --- | --- |
| System Time | Current North American Central Time Zone |
| Operation Mode | Manual |
| Programming Step Type | Non-ramping |
| Temperature Set Point | 25°C |
| Temperature Limit 1 High Alarm | 35°C |
| Temperature Limit 1 Low Alarm | -5°C |
| Temperature Limit 2 High Alarm | 38°C |
| Temperature Limit 2 Low Alarm | -8°C |
| Light Output 1 | 0% (Off) |

In general, any optional features controlled via the IntellusUltra are disabled by default.  Please refer to the user manuals provided with optional equipment for information pertaining to their factory settings, setup and operation.

### Calibration

All controlled environment chambers from Percival Scientific must pass rigorous quality checks and validation testing prior to shipment.  As part of the validation process, all chamber control systems are calibrated against an NIST standard device at the following conditions:

| Condition | Calibration |
| --- | --- |
| Temperature w/ Lights Off | 25°C |
| Temperature w/ Lights On (for lighted models) | 25°C |
| Relative Humidity w/ Lights Off (if RH option is equipped) | 75%RH @ 25°C |
| Relative Humidity w/ Lights On (if RH option is equipped) | 75%RH @ 25°C |

At these calibration points, any discrepancy between an IntellusUltra sensor (located by the intake of the evaporator coil) reading and the NIST calibration standard (located in the center of the chamber environment) are entered into the control system as calibration offsets.

For any given environmental parameter controlled by the chamber:

Calibration Offset=Calibration Standard Value-Control System Process Value

For example, if the control system reads a temperature process value of 25°C and the calibration standard reads 24.5°C, then an offset of -0.5°C is saved into the controller.  For each condition, the measured offset is applied to the process value so that it matches the calibration standard.

TOffset= TStandard- TProcess Value

TOffset= 24.5°C- 25.0°C

TOffset= -0.5°C

Any offsets with lights turned off are called Night Offsets and any offsets with any amount of lights turned on are called Day Offsets.

Please refer to the provided control system user manual for additional information regarding calibration offsets.

Factory calibration offsets serve as suitable values for general chamber use.  Consider recalibrating the chamber if precise environmental control is required at specific conditions that are significantly different than the factory calibration points.

# Operation

## Guidelines for Loading a Chamber

When placing products, specimens or experiments onto bench flooring inside a PGC-15 controlled environment, adhere to the following guidelines to ensure the unit operates at peak performance:

- For best temperature uniformity, evenly distribute items throughout the chamber space and allow space between items to promote proper air flow

- As a good rule of thumb, do not populate more than 80% of the bench flooring’s surface area

- Do not place more than 500 lbs (227.3 kg) total in a chamber

Avoid placing items in a product that produce significant amounts of heat.  PGC-15 chambers are designed to accommodate a product heat load of 2 W/ft² of shelf space.  Any heat introduced into the chamber environment beyond this amount can result in loss of environmental control.

## Fresh Air Exchange

Adjustable fresh air ports are proved as standard equipment on the PGC-15 controlled environment chamber.  These ports are located below the main chamber doors and are designed to allow an air exchange between the interior and exterior of the chamber. They are configured to prohibit light from passing through the ports.

The fresh air port located closest to the chamber’s mechanical end section is the exhaust, and the port located farthest from the mechanical end section is the intake.

Intake

Exhaust

Please note that fresh air ports are less insulated than the walls into which they are installed and are prone to the formation of condensation on their external surfaces.  Also, opening the ports introduces ambient air into the chamber which can work against humidity and temperature set points inside the controlled environment.

## Running a Control Profile

- Close the main chamber door, ensuring a proper seal between the door gasket and insulated box

- Configure static or multi-step profile process value set points

- Check alarm limits, ensuring that they are set above the maximum and below the minimum process values.  Limits should not be set closer than 2-3°C above or below the operational range of the entire control profile.

- Select the mode of operation on the control system (e.g. Manual, Diurnal, Program or Sequence).

Once an operating mode is set, the IntellusUltra control system automatically executes any set points specified in the profile.

If an alarm threshold is exceeded, the controller produces an audible alarm until the alarm is acknowledged by the user.  The alarm state remains active until the source of the alarm is resolved.  A temperature alarm condition shuts down all chamber sub-systems inside the environment until the alarm state is cleared.  Whenever process values fall back within allowable limits, the control system automatically reverts to normal operation, requiring no involvement from the user.

The IntellusUltra controller has an internal battery which allows it to keep time in the event of a power cycle or outage.  Upon power up, the chamber will resume the previously executed control profile.

- Control is resumed around static process value set points when the controller is in manual mode

- When executing a multi-step program, the controller resumes control around the current program step as determined by the controller system time.

- When executing a sequence of programs, the controller resumes control around the current program iteration and step as determined by the system time unless power is off at 12:00 am.  If power is off at 12:00am the program iteration count will not be incremented.

For best control, minimize the frequency and duration of chamber door openings.

# Maintenance

Performing regular maintenance on an environmental chamber reduces performance degradation, improves reliability and extends operational life.  In addition to any user-specific maintenance protocols, the following general maintenance tasks should be performed at the recommended intervals.

Turn chamber power off prior to performing any maintenance inside the environment.

Unplug or disconnect chamber supply power prior to performing any maintenance inside the chamber’s mechanical compartment or control box.

## General Maintenance Summary Table

| Maintenance Task | Recommended Interval Schedule |
| --- | --- |
| Clean the Chamber | Every 3 months, or as needed |
| Clean the Bench Flooring/Air Diverters | Every 3 months, or as needed |
| Clean the Floor/Floor Drain | Every 3 months, or as needed |
| Clean the Condenser | Every 3 months, or as needed |
| Clean the Evaporator | Every 12 months |

## Cleaning the Chamber

Wipe down all interior and exterior surfaces with a solution of water and mild detergent.  Do not use abrasive agents or cleaning pads as these could scratch surfaces.  Glass cleaners can be used on stainless steel surfaces to prevent streaking.

## Clean the Bench Flooring & Air Diverters

Remove bench flooring from chamber and clean by brushing with solution of water and mild detergent, vacuuming and/or using compressed air.

When re-installing bench flooring sections, please note that the air flow perforations are different on each side of the bench sections.  For proper air flow, position each bench section such that the side with more air flow holes are pointed toward the chamber’s door openings.

Air diverter panels

High-density hole pattern toward front

Remove and clean

bench flooring

After removing bench flooring, air diverter panels will be visible.  Remove the cover panels beneath the perforated bench to access the evaporator fans, evaporator coils(s), heater(s), and floor drain:

## Clean the Evaporator Coil(s)

Once the unit is powered down, remove any dust or debris collected upon the evaporator fins, copper tubing, and impeller blades.  Clean the evaporator coil by dislodging dirt and debris via soft brush, compressed air or vacuum.  Wipe dust and debris from the impeller blades.

If the evaporator has a brown phenolic coating, please refer to the supplemental Q9 Phenolic Coating User Manual for specific cleaning requirements in lieu of these general guidelines.

## Clean the Floor & Floor Drain

After removing the air diverter panels, the chamber interior floor & floor drain will be visible.  Clean by brushing with solution of water and mild detergent.

Clean floor &

floor drain

## Clean the Heat Exchangers

Air must pass freely through the evaporator and condenser for proper performance.  Ambient air is drawn into and forced through the condenser by the condensing unit fan.  Chamber process air is drawn into and forced through the evaporator by a bank of fans above the evaporator.  In both cases, the fans run continuously while the chamber is in operation and there are no active alarms.

Power down the chamber prior to any cleaning effort on refrigeration system components.

Please be aware that some parts of the condensing unit become hot while the unit is in operation and can remain hot for some time after the unit is powered down.

### Clean the Condenser

Once the unit is powered down, remove any dust or debris collected upon the condenser fins, copper tubing, fan blades and motor.  Blow dust out of the condenser from behind with compressed air, a brush, or vacuum and wipe dust from the fan blades.

To access the condenser coil within the lower mechanical section, unscrew and remove the vented front service panel underneath the control panel.

Front service panel

Clean condenser fins

To access the condenser fan, unscrew and remove the lower mechanical section’s side cover.

Remove fan guard & clean blades & motor

Side cover

### Clean the Aspirator Box Fan

Once the unit is powered down, remove any dust or debris collected upon the aspirator fan blades and motor.  Remove box cover and clean the aspirator fan by wiping dust and debris from the fan blades and motor.

Box cover

# Troubleshooting

## Loss of Temperature Control – General

| Possible Reason | Solution |
| --- | --- |
| Improper or incorrect signal between controller and control relay(s) | Check for ~5VDC output signal between the controller and solid-state relay. Output state is indicated by LEDs on the controller overlay. Contact factory if insufficient output voltage from the control system |
| Temperature sensor | Check the sensor for correct output signal. Refer to the RTD temperature vs. resistance table in the controller manual. Contact factory if sensor output does not match the table values. |
| Solid-state relays | If correct trigger voltage is correct to the relay coil, check for the correct output voltage through the relay. Refer to the chamber’s electrical schematic for the VAC output to the affected heating and/or cooling device(s) |
| Hot gas solenoid valve | When in heating mode, the hot gas solenoid should be in the open position allowing hot gas to flow into the evaporator coil. Check the solenoid valve for proper actuation and the valve body for internal blockages. |
| Liquid line solenoid valve | When in cooling mode, the liquid line solenoid valve should be in the open position, allowing liquid refrigerant to flow into the thermal expansion device(s). Check the solenoid for proper actuation and the valve body for internal blockages. |

## Loss of Temperature Control - Overheating

| Loss of refrigerant | Check system pressures and check for leaks with a leak detector. Evacuate the system, repair leak(s) and recharge the system. Refer to the unit’s refrigeration diagram and serial label. |
| --- | --- |
| Restriction in the refrigeration piping or componentry | If adequate pressure and no leaks, there may be a restriction in the unit’s liquid line solenoid valve, filter/drier or thermal expansion device. Check the line temperature directly before and after each device. There may be a blockage if the outlet is colder than the inlet to the solenoid and drier. |
| Insufficient heat exchange through the condenser | Check the ambient air temperature is below 90°F. Make sure the condenser fan and coils are clean. Make sure the condenser fan is running. |
| Compressor motor fails to pump refrigerant throughout the system | Check electrical power to the motor and start components. Perform an ohm check on the motor windings. Check high-side system pressure if the unit runs |
| Insufficient heat exchange through the evaporator | Check for proper fan operation. Fans should be continuously running, drawing air into the evaporator. Make sure the evaporator coils are clean. |

## Loss of Temperature Control – Overcooling

| Electric heater malfunction | Check output voltage from the controller to the heater relay. Check output voltage from the heater relay to the heater element. Check the circuit breaker in the heater electrical circuit. |
| --- | --- |
| Malfunctioning hot gas solid-state relay (capillary tube systems) | Check output voltage from the controller to the hot gas line solenoid valve. Check voltage between the relay and solenoid valve. The solenoid valve should be open whenever the ‘heat’ LED is lit |
| Malfunctioning liquid solid-state relay (TXV systems) | Check output voltage between controller and relay. Check volage between relay and solenoid. The solenoid should fully close when the ‘heat’ LED indicator light is energized. |
| Blocked hot gas solenoid valve (capillary tube systems) | If the hot gas solenoid valve gets blocked, then flow will divert through the capillary tube, causing the system to cool even when calling for heat. |

## General

| Chamber does not power on | Check facility electrical services and circuit breakers. Check that the chamber is properly wired into the facility services. Check circuit breakers inside the unit’s control box. |
| --- | --- |
| Individual chamber component does not turn on | Check controller settings and verify that the device has been properly enabled. Check the control box for tripped circuit breakers. |
| Condensation | The amount of condensation that can form inside and be rejected from a controlled environment can fluctuate, and is highly dependent upon ambient conditions, interior environmental conditions, frequency of door openings, product load and optional humidity control systems. Minimize door openings and the amount of moisture present inside the unit, and close the access port |

## Parts and Service

Replacement parts can be ordered from Percival Scientific at:

Ph: 1-800-695-2743

Email: parts@percival-scientific.com

For additional questions regarding this product, please contact Percival’s customer support team at:

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com