Walk-in Chamber

Operator’s Manual

505 Research Drive  Perry, Iowa 50220

Toll-free 800.695.2743  515.465.9363  fax 515.465.9464

www.percival-scientific.com

[This page intentionally left blank]

# Introduction

To help customers create better science, Percival Scientific continuously studies and refines the design of its products to provide the best possible solutions and equipment in the industry.  As such, changes in the design and/or operation of equipment may occur without publication.  Please feel free to contact Percival Scientific directly for product updates.

Please read this manual carefully before proceeding with operation

# Revision Log

| Version | Date | Description |
| --- | --- | --- |
| v0 | 7/22/2025 | Draft |
| R1 | 7/25/2025 | Initial release |
| R2 | 10/31/2025 | Added FR information; modified Access Port information |

Introduction	3

Revision Log	4

Warranty Information	8

Definitions	10

General	10

Equipment Overview:	11

Installation	12

Site Preparation	12

Location	12

Ambient Conditions	13

Chamber Heat Rejection to Ambient	13

Electrical Services	13

Drainage and Plumbing Services	14

Unpackaging	15

Inspect the Shipment	15

Remove Chamber Components from Skids	16

Installation Guidelines	16

Connect Services	16

Startup and Commissioning	17

Inspect the Control Box	17

Inspecting Electrical Connections	17

Turn on the Chamber	18

IntellusUltra Control System	18

Calibration	20

Operation	21

Guidelines for Loading a Chamber	21

Running a Control Profile	21

Convenience Receptacle Operation	22

Lighting	23

Air Flow	24

Horizontal Air Flow	24

Perforated Plenums	25

Vertical Upward Air Flow	26

Vertical Downward	26

Mixed Air Flow	26

Air Flow Summary Table	27

Access Ports	27

Size and Location	27

Installation	27

Operation	28

Condensation	28

Fresh Air Exchange	28

Intake	29

Exhaust	29

Exchange Rate	29

Operation	29

Condensation	30

Sensor Aspirator Box	30

Dry Contacts	32

Maintenance	32

General Maintenance Summary Table	32

Cleaning the Chamber	33

Clean the Heat Exchangers	33

Clean the Condenser (Air Cooled)	33

Clean the Evaporator	34

Inspect the Aspirator Box Assembly	35

Troubleshooting	36

Loss of Temperature Control – General	36

Loss of Temperature Control - Overheating	36

Loss of Temperature Control – Overcooling	37

General	37

Parts and Service	38

# Warranty Information

Equipment Limited Warranty

Thank you for purchasing equipment from Percival Scientific, Inc.  Our goal is to provide every customer with an exceptional ownership experience.  We are committed to providing products and services of the highest quality, and your satisfaction is our top priority.  In support of this objective, we present the following warranty.

Warranty Terms

Percival Scientific, Inc. (hereinafter referred to as Percival) warrants that equipment manufactured by Percival shall be free from defect in material or workmanship, when installed, operated and maintained in accordance with instructions provided by Percival.

Warranty Coverage

- Coverage period shall be twenty-four (24) months from the date of delivery

- Percival’s obligation under this warranty is limited to the repair or replacement of defect at no charge for parts, shipping and labor

- Percival reserves the right to inspect the equipment after a warranty claim is made to determine the proper course of corrective action

Exclusions

- Parts or elements damaged from improper installation, including mechanical, electrical and plumbing services, IT infrastructure, inadequate room ventilation, improper water quality, and ambient conditions that are out of specification

- Parts or elements damaged from improper operation

- Parts or elements damaged from improper maintenance

- Parts or elements considered to be consumable, that outlive their normal expected terms of usefulness, such as: light bulbs, filters, desiccant material and CO2 scrubbing media)

- Damage resulting from shipment to or relocation within the customer’s facility after equipment has been signed for and ownership is transferred to the customer

- Modification to equipment by the customer without Percival’s consent may void this warranty

- Repair work undertaken without prior consent of Percival may void this warranty

- Damage caused by accident, fire, Force Majeure or Acts of God

Additional Terms

- This warranty is in lieu of any other warranties, expressed or implied, including merchantability or fitness for a particular purpose

- Warranty terms described herein cannot be changed without the expressed, written consent from Percival Scientific, Inc.

- The owner agrees that Percival’s sole liability with respect to defective parts or materials shall be set forth in this warranty, and any claims for incidental or consequential damages are expressly excluded

- This warranty is non-transferable and applies only to the original purchaser

- Repaired or replaced parts shall be warranted only for the unexpired portion of the original warranty, unless otherwise defined

- Parts replaced under warranty shall become the property of Percival

To Make a Warranty Claim

- Document the issue with detailed notes and photographs along with the equipment’s serial number

- Contact Percival’s customer support team at customerservice@percival-scientific.com, parts@percival-scientific.com or 1-800-695-2743

- Once your claim is received, our customer support team will review it and contact you with further instructions

# Definitions

‘Controlled environment chamber’, ‘Chamber’, ‘Walk-in’ and ‘Unit’ are all used interchangeably in this manual to describe the complete controlled environment chamber system manufactured by Percival Scientific, Inc.

# General

This manual is intended to provide instructions on the set-up, operation, maintenance and troubleshooting of Percival Scientific’s walk-in chamber product lines, including AR, CTH, FR, IR, MPR, PR, and SS models.

Walk-in chambers play a crucial role in various scientific fields of study, serving as controlled environments for the growth and maintenance of a wide range of biological specimens. They are highly valued in academic, government, and commercial research laboratories where environmental accuracy and consistency are critical for experimental repeatability.  These precision-controlled environments designed for a wide variety of large-scale scientific research applications, enable reproducible experimental conditions by regulating temperature, relative humidity, photoperiod, and light intensity. They are suitable for both short- and long-term studies.

This manual outlines the procedures and best practices for operating and maintaining Percival Scientific’s various walk-in chamber base-models.  Any special features not covered in this user manual shall be described in a supplemental information section.

Users should read this manual thoroughly to ensure proper setup, calibration, programming, and maintenance to avoid damage to specimens or equipment and to ensure data integrity.

## Equipment Overview:

(PR Series w/ SWR Lighting)

# Installation

To achieve and maintain optimum performance from the unit throughout its designed operating range(s), it is important to install the chamber according to the following guidelines.  Failure to follow the installation guidelines in this section may cause loss of environmental control, malfunction, premature system failure and may also void the manufacturer’s warranty.

In most cases, walk-in chamber installations are performed by Percival’s field installation team in coordination with local contractors, authorities having jurisdiction (AHJ), project stakeholders and end users.  When this is the case, the following aspects are taken into consideration prior to and during the installation process.

## Site Preparation

The following considerations should be made when preparing a facility for installation of Percival controlled environments:

### Location

Locate the chamber in a clean area with a level surface and adequate space for door swings, ventilation, access to facility services and the unit’s service connections and access panels.

Prior to the unit’s delivery date, ensure there is a clear path from the delivery dock to final chamber location, taking doorways, elevators, corridor widths and turns into account.

Warning! Walk-in components and subassemblies can be heavy.  Use caution when removing chamber components from their shipping containers and while moving them into position.  These tasks should be undertaken with care by more than one person.

Spatial requirements vary by chamber model and optional equipment.  Please refer to all supplementary installation instructions as optional equipment can change the space required for a chamber’s proper operation.

Generally, space should be planned for door swings, removal of service panels, condensing unit intake and exhaust ventilation and shutoff valves for chamber services.

Provide at least two inches of clearance around the walk-in’s insulated wall panels, and at least six inches of clearance around intake and exhaust ports to facilitate proper air flow through into and out of the chamber.

Additional space should be planned around the chamber’s refrigeration system, such that proper ventilation is available for air-cooled units and sufficient space is provided for service and maintenance purposes.

### Ambient Conditions

Published chamber performance specifications are based on the following ambient environmental conditions.  To ensure best performance of a controlled environment chamber, locate the unit in an environment where the following conditions can be maintained:

Temperature = 75 °F (23.8 °C)

Relative Humidity = 55 %RH

### Chamber Heat Rejection to Ambient

To achieve and maintain specified conditions within a controlled environment chamber, the unit’s cooling system must remove excess heat from within the chamber.  This excess heat is rejected from the system and must be accounted for when planning for a chamber’s installation.

It is necessary to ensure the HVAC system where the chamber is installed has sufficient capacity to remove heat rejected by all chambers within the space.  Insufficient heat removal capacity will result in elevated ambient temperatures and possible degradation in chamber performance.

Each walk-in chamber model’s estimated maximum heat rejection to ambient can be found in the Product Specifications, or Supplemental Information section of this manual.

### Electrical Services

A walk-in controlled environment chamber is typically provided with terminal blocks inside its control box that must be connected to electrical building services of appropriate type, rating and ampacity.  Final electrical connections are often coordinated through a local electrician.

Always verify a chamber’s electrical requirements against the provided electrical services prior to making final connections and energizing the unit.

A chamber’s electrical requirements are listed on its serial tag, which is located inside the unit’s control panel.  These electrical requirements include the operating voltage, phase(s) and maximum running load amperage (RLA) for the chamber, per the following requirements:

- Supply voltage must be within ±10% of the serial tag value.

- The number of phases provided should match the serial tag value.

- The electrical service must have a minimum circuit ampacity (MCA) sufficient to support the chamber’s rated maximum running load amperage (RLA).  Service MCA should ultimately be determined by local authorities having jurisdiction (AHJ), according to state, local and building electrical codes.  As a rule of thumb, MCA can be estimated by the following equation:

MCA=1.25×RLAmax

- The provided electrical service should not be smaller than the MCA.  In many cases the service itself will be larger than the MCA. For example, a standard 30A electrical service might be specified to comply with a chamber’s MCA requirement of 27A.

If a chamber is provided with multiple electrical service connections, each connection will have its own RLA/MCA requirement.  If all connections are to be supplied through a single service, the service must be sized for the COMBINED RLA of all connections.  An example equation is provided below for reference:

MCAcombined=1.25 ×(RLA 1+ RLA 2)

In lieu of a single electrical service sized for multiple devices, connect each device to its own separate, dedicated electrical service.

Failure to follow the electrical service requirements outlined above can result in overloaded electrical services, tripped circuit protection on the supply and damage to facility and chamber electrical infrastructure such as wiring, plugs and receptacles.

Damage caused from incorrect or improperly sized electrical services is not covered under the manufacturer’s warranty.

### Drainage and Plumbing Services

Walk-in controlled environments with refrigeration-based cooling systems reject condensed water from their evaporator coils.  The amount of condensate is dependent upon a combination of controlled environment settings, ambient conditions, door opening frequency and duration, fresh air exchange for occupancy, and characteristics of products placed inside the environment.

Optional relative humidity control systems increase the amount of condensate rejected from a given unit.

Whenever the evaporator coils operate at a temperature below the dew point inside the controlled environment, moisture present in the circulated air condenses onto the cooling coils.  This condensate sheds off the coils, is routed through a drain and expelled out of the environment’s floor drain.

Base model walk-in chambers utilize a single evaporator coil assembly with one condensate drain.  Ideally, this drain line should be routed toward a floor drain inside the environment, or in close proximity to the chamber to properly manage condensate rejected from the unit.

If the installation location for the chamber does not provide a nearby floor drain, the following drain accommodations can be considered:

- Extend the existing chamber drain line external to the unit and route it to a floor drain that is farther away than the chamber’s footprint

- Connect the existing chamber drain system to a condensate pump

Please note that the standard drain system in Percival’s walk-in controlled environment chambers work via gravity.  Condensate only drains properly if the facility drain is lower in elevation than the chamber drain line.  Proper slope for condensate drain lines is at least 0.125 in of fall towards the facility drain for every 12 in of horizontal run.

## Unpackaging

Unless otherwise defined, shipping logistics and materials handling for all walk-in chamber components are included when field installation services are performed by Percival technicians.  When chamber installation is handled by other parties, the following considerations for receiving and unpackaging the unit should be made:

### Inspect the Shipment

At the time of delivery, it is very important to verify that the proper quantity and type of products were received as per the Bill of Lading, and to inspect those products for shipping damage prior to accepting the shipment.

- Carefully remove all packaging materials that surround the unit, ensuring no loose chamber components are accidentally discarded.

- Visually inspect both the interior and exterior of the unit for dents, scratches or any other signs of damage

- It is extremely important to note any damage on the shipping receipt so that corrective actions can be taken as efficiently and expeditiously as possible.  The shipping documentation provides the following options:

  - Accept the shipment as-is

  - Accept the shipment with damages documented, understanding that corrective actions to rectify the damage will be performed onsite

  - Reject the shipment, where damage is sufficient to require the product to be returned to Percival, where corrective actions can be taken at the manufacturing facility

### Remove Chamber Components from Skids

Warning! Walk-in chamber components are heavy.  Please use caution when removing large parts from their crating and while moving them into position.  This task should be undertaken with care by more than one person.

## Installation Guidelines

Refer to supplemental installation instructions and photos provided by Percival, when installation of a walk-in chamber is to be performed by a third party.

### Connect Services

- Connect properly sized electrical services to the chamber control box and any other service connection points.  Due to its large electrical demand and possible remote location from the chamber, the condensing unit is often specified with a dedicated electrical service separate from the main chamber electrical service.

- Route the chamber’s drain to a floor drain or other drainage accommodation

- Connect services for optional equipment as directed in supplemental operation manuals (e.g. demineralized water, compressed gases, chilled water, ventilation ductwork, etc...)

## Startup and Commissioning

### Inspect the Control Box

While unplugged from electrical services, open the chamber’s control panel and perform a visual inspection of the electrical components installed within it.

- Check the state of the resettable circuit breakers and reset any found in an ‘open’ state

- Check for any components that are not physically secured to the components mounting plate, control box housing or top of the insulated box.

- Check for any connector plugs that are dislodged or not fully seated into their receptacles and re-secure them as needed with a firm press.

- Check for any obviously loose or disconnected wiring

If any components have become dislodged or there is any question as to where a disconnected plug or wire should be connected, do not proceed any further with chamber installation or operation and contact Percival’s customer support team at 1-800-695-2743 or customerservice@percival-scientific.com

### Inspecting Electrical Connections

Make sure all plugs are fully seated into their proper receptacles with a firm press.  Failure to fully seat a plug into its receptacle may cause the powered device and/or electrical circuit to malfunction or become damaged.  LED lamp banks are particularly sensitive to plug misalignment and insufficient engagement between plug(s) and receptacle(s).  An example of this is provided below for reference:

(Plug not fully seated)		             (Press plug firmly into receptacle)

### Turn on the Chamber

Set the main power switch located on the left side of the control panel to the ‘On’ or ‘|’ position to energize the chamber.  The IntellusUltra controller takes approximately 15-20 seconds to start up.  Once the controller has started up and there are no active alarm conditions, the rest of the chamber components will energize.

### IntellusUltra Control System

The primary method of operating a walk-in controlled environment chamber is through the C9T IntellusUltra control system’s touch screen. The touch screen (shown below) is located on the exterior of the walk-in base model’s control panel.

Please refer to the supplementary Touch Screen User Manual for instructions on the operation of this system

Refer to the IntellusUltra Quick Guide provided with the chamber for basic instructions on the controller’s operation.  This guide can also be downloaded at https://www.percival-scientific.com/wp-content/uploads/2023/03/IntellusUltra-Quick-Guide.pdf

Refer to the C8/9 Control System User Manual for comprehensive operational instructions for the chamber’s control system

When first energized, the IntellusUltra control system is set to operate in Manual Mode with the following factory settings:

| Control Feature | Factory Setting |
| --- | --- |
| System Time | Current North American Central Time Zone |
| Operation Mode | Manual |
| Programming Step Type | Non-ramping |
| Temperature Set Point | 25°C |
| Temperature Limit 1 High Alarm | 35°C |
| Temperature Limit 1 Low Alarm | -5°C |
| Temperature Limit 2 High Alarm | 38°C |
| Temperature Limit 2 Low Alarm | -8°C |
| Light Output 1 (LL, VL and LLVL models) | 0% (Off) |
| Light Output 2 (LLVL models) | 0% (Off) |

In general, any optional features controlled via the IntellusUltra are disabled by default.  Please refer to the user manuals provided with optional equipment for information pertaining to their factory settings, setup and operation.

Note: The keypad overlay (shown below) inside the control box is used for service purposes

### Calibration

All controlled environment chambers from Percival Scientific must pass rigorous quality checks and validation testing prior to shipment.  As part of the validation process, all chamber control systems are calibrated against an NIST standard device at the following conditions:

| Condition | Calibration |
| --- | --- |
| Temperature w/ Lights Off | 25°C |
| Temperature w/ Lights On (for lighted models) | 25°C |
| Relative Humidity w/ Lights Off (if RH option is equipped) | 75%RH @ 25°C |
| Relative Humidity w/ Lights On (if RH option is equipped) | 75%RH @ 25°C |

At these calibration points, any discrepancy between an IntellusUltra sensor (located by the intake of the evaporator coil) reading and the NIST calibration standard (located in the center of the chamber environment) are entered into the control system as calibration offsets.

For any given environmental parameter controlled by the chamber:

Calibration Offset=Calibration Standard Value-Control System Process Value

For example, if the control system reads a temperature process value of 25°C and the calibration standard reads 24.5°C, then an offset of -0.5°C is saved into the controller.  For each condition, the measured offset is applied to the process value so that it matches the calibration standard.

TOffset= TStandard- TProcess Value

TOffset= 24.5°C- 25.0°C

TOffset= -0.5°C

Any offsets with lights turned off are called Night Offsets and any offsets with any amount of lights turned on are called Day Offsets.

Please refer to the provided control system user manual for additional information regarding calibration offsets.

Factory calibration offsets serve as suitable values for general chamber use.  Consider recalibrating the chamber if precise environmental control is required at specific conditions that are significantly different than the factory calibration points.

# Operation

## Guidelines for Loading a Chamber

When placing products, specimens or experiments onto shelves inside a controlled environment, adhere to the following guidelines to ensure the unit operates at peak performance:

- For best temperature uniformity, evenly distribute items throughout the chamber space and allow space between items to promote proper air flow

- As a good rule of thumb, do not populate more than 80% of any shelf’s surface area

- Do not place more weight on a shelf than its static loadbearing capacity rating

Avoid placing items that produce significant amounts of heat in a unit.  Percival’s walk-in chambers are designed to accommodate a specific product heat load of 2-3 W/ft² of shelf space, depending upon the specific model number.  Any heat introduced into the chamber environment beyond this amount can result in loss of environmental control.

## Running a Control Profile

- Close the main chamber door, ensuring a proper seal between the door gasket and insulated box

- Configure static or multi-step profile process value set points

- Check alarm limits, ensuring that they are set above the maximum and below the minimum process values.  Limits should not be set closer than 2-3°C above or below the operational range of the entire control profile.

- Select the mode of operation on the control system (e.g. Manual, Diurnal, Program or Sequence).

Once an operating mode is set, the IntellusUltra control system automatically executes any set points specified in the profile.

If an alarm threshold is exceeded, the controller produces an audible alarm until the alarm is acknowledged by the user.  The alarm state remains active until the source of the alarm is resolved.  A temperature alarm condition shuts down all chamber sub-systems inside the environment until the alarm state is cleared.  Whenever process values fall back within allowable limits, the control system automatically reverts to normal operation, requiring no involvement from the user.

The IntellusUltra controller has an internal battery which allows it to keep time in the event of a power cycle or outage.  Upon power up, the chamber will resume the previously executed control profile.

- Control is resumed around static process value set points when the controller is in manual mode

- When executing a multi-step program, the controller resumes control around the current program step as determined by the controller system time.

- When executing a sequence of programs, the controller resumes control around the current program iteration and step as determined by the system time unless power is off at 12:00 am.  If power is off at 12:00am the program iteration count will not be incremented.

For best control, minimize the frequency and duration of chamber door openings.

## Convenience Receptacle Operation

Convenience receptacles are often provided in walk-in chambers for powering auxiliary equipment.  Electrical loads generate significant amounts of heat, which can be problematic when placed inside of a controlled environment.  To optimize system performance, convenience receptacle circuit capacity may be limited by branch circuit protection to amperages that are lower than the maximum ratings of the circuit’s infrastructure.

For example, multiple duplex receptacles provided inside a walk-in may be wired into a single circuit and limited to 5A total, even though the individual duplex receptacles may be rated for up to 15A or 20A.

In addition to limits applied to the circuit’s capacity, the duty cycle of convenience receptacles must also be considered.  Unless specified, convenience receptacle circuits should not be continuously energized as the heat wattage introduced into the controlled environment can place a significant demand upon the unit’s refrigeration system.  Downstream negative effects can include:

- Inability to achieve specified process values

  - low temperature set points

  - low and high relative humidity set points

- Slow system responsiveness

- Inability to maintain specified process value control stability

- Large temperature and relative humidity gradients across the experiment space

- Loss of temperature control resulting from ice buildup on the evaporator coil(s)

## Lighting

The standard lighting system provided on a walk-in is specific to each chamber model type and any customer requirements defined in project requirement documentation.  The standard lighting system included with each chamber model is provided in the table below:

| Chamber Model | Standard Lighting System |
| --- | --- |
| AR | SWR (SciWhite® with Red) |
| CTH | IW (IncuWhite®) |
| FR | IW (IncuWhite®) |
| IR | IW (IncuWhite®) |
| MPR | SWR (SciWhite® with Red) |
| PR | SWR (SciWhite® with Red) |
| SS | IW (IncuWhite®) |

Please refer to the supplemental lighting system user manual provided with this document that corresponds with the lighting system option code ordered with the chamber.  Non-standard lighting system options may include one or more of the following: SB (SciBrite®), SW (SciWhite®), IW (IncuWhite®), IB (IncuBrite™), PC (PetriClear®), or X (Special)

## Air Flow

Air flow is tuned to provide sufficient capacity to remove excess heat from within the chamber and establish proper uniformity while minimizing disturbance to specimens.

Air is continuously recirculated inside Percival’s walk-in units, where it is conditioned and distributed throughout the environment in the following manner for each model type.  Lighting and shelf construction are the major determining factors regarding air flow management inside a walk-in chamber.  There are four basic flow patterns, when each chamber model utilizes one specific pattern for optimal performance.

### Horizontal Air Flow

This air flow pattern is necessary on chambers with multiple tiers, solid over-shelf lighting, and/or solid shelving.  The chamber system is designed to move process air horizontally across each tier via adjustable plenums affixed to the side walls of the unit.

(Generic AR / MPR model shown above)

### Perforated Plenums

Perforated plenums are used to develop uniform, horizontal air flow across each tier.  The plenum affords some adjustment of the air flow amount across each tier.  This is accomplished through an adjustable panel that can be used to restrict flow through each plenum section.

Temperature differences between shelf tiers can be minimized through adjustment of the plenums.  As a rule of thumb, adjusting a plenum section open will increase the rate of air flow, and decrease the shelf temperature relative to the other tiers.  Adjusting a plenum to a closed position will have the reverse effect.  When adjusting plenums, please keep in mind that there is a fixed total amount of air movement capacity in the chamber system. Adjustment of one plenum section will either steal air from or divert excess air toward the other plenum sections.

Do not remove the screws or the air block panel will detach.

To adjust a plenum section, loosen the adjustment screws only enough to allow the air block panel to slide horizontally.  Slide a plenum adjustment panel left to increase air flow and right to reduce air flow.  Once the panel has been adjusted, retighten the screws.

Air Block

Adjustment Screw

Air Block

Adjustment Screw

### Vertical Upward Air Flow

Upward air flow is utilized on chambers with adjustable-height, counterbalanced lamp canopies.  These lighting systems typically provide high light intensities, and the vertical upward air flow directs unwanted lamp heat directly into the evaporator coils for removal from the chamber environment.

(Generic PR model shown above)

### Vertical Downward

A vertical downward air flow pattern is essentially opposite of upward flow, as described above.  Upon request, walk-in chamber models with upward air flow can typically have their airflow pattern reversed to downward.

### Mixed Air Flow

Percival’s CTH, FR, IR and SS walk-in models work well with wire shelving and low-profile light fixturing.  These elements allow for process air to flow throughout the environment without the need for discreet, directional air flow patterns.

In a mixed air flow pattern, process air is drawn upward at the center of a chamber into the ceiling-mounted evaporator(s) where it’s conditioned.  Process air is then expelled from the evaporator(s) toward the chamber’s side walls, where it cascades down into the experiment space.  Open wire shelving and low-profile lighting allow process air to flow both vertically and horizontally over experiments before returning to the evaporator(s) for reconditioning.

[insert CTH/IR air flow diagram here]

### Air Flow Summary Table

| Chamber Model | Air Flow Direction |
| --- | --- |
| AR | Horizontal |
| CTH | Mixed |
| FR | Mixed |
| IR | Mixed |
| MPR | Horizontal |
| PR | Upward |
| SS | Mixed |

## Access Ports

Access ports are designed to provide a pathway between a chamber’s interior and its surrounding environment.  These ports are commonly used to pass instrumentation, cabling or tubing into an environment, while allowing its door to remain closed.

AR, CTH, IR, MPR and PR walk-in chamber models from Percival Scientific can be configured with access ports.

### Size and Location

The diameter and location of an access port(s) is specific to each chamber model and is typically defined by the customer during the ‘Discovery and Planning’ phase of Percival’s Proven Process.

Unless otherwise specified, the standard-sized access port has an internal pass-through diameter of 1 ½” (38mm).  Alternative pass-through sizes include 2”, 3” and 4” diameters, and custom sizes are also available upon request at the time of order.

### Installation

Each access port assembly is provided with one (1) pass-through sleeve (or port) and one (1) insertion plug or two (2) screw-on caps.

No tools are required for installation of port plugs or caps.

Plugs and caps can be found inside the accessories bag provided with each chamber.  Ports are left open during shipping to allow some fresh air exchange and to prevent stagnation inside the environment.

To install a port plug, insert it into its port sleeve from outside of the chamber.  The sleeve should fit snugly into its sleeve.  Fit can be adjusted by tightening (clockwise) or loosening (counterclockwise) the wingnut on the port plug.

If the access port system utilizes caps instead of plugs, place the cap over one of the threaded ends of the access port sleeve and tighten it down with a clockwise turn until finger tight.  Note: Ports with caps are most typically used on chambers with walls thicker than 2” (50.8mm).  This style of port has an internal and an external cap.

### Operation

When access ports are in use (i.e. plugs or caps removed and implements inserted), it is recommended that any remaining voids within the port are sealed to prevent ingress of ambient air into the controlled environment.  Packing foam or fabric into a void or sealing externally with tape are all effective means to minimize disturbance to the controlled environment.

When not in use, the plug (or cap) provided with each port can be installed to create an air-tight seal between the chamber and its surrounding environment.

### Condensation

Access ports are generally less insulated than the chamber walls into which they are installed and are thus prone to the formation of condensation on their internal or external surfaces.  The degree to which condensation may form is dependent upon conditions inside the controlled environment relative to ambient.

When opened (plug or caps removed), each access port allows some unmanaged air exchange between the controlled environment and ambient environment which can further the development of condensation.  Generally, this condensation is minor and can be resolved by the addition of insulative material inside the port sleeve and/or measures to seal off any air exchange.

## Fresh Air Exchange

Unless otherwise specified, Percival’s walk-in models are provided with a fresh air exchange system, which is necessary to support chamber occupancy requirements for research personnel.  Additionally, the fresh air exchange system facilitates diffusion and respiration processes in biological specimens.

### Intake

The fresh air intake port is located close in proximity to the evaporator(s) so that ambient air drawn into the chamber can be immediately conditioned, minimizing disturbance to the controlled environment.  Evaporator fans help evenly distribute conditioned fresh air throughout the chamber environment.  The port includes sheet metal covers which control the direction of flow and prevent light contamination into and out of the environment.  The intake port can be closed with an internal slide cover to restrict the flow of ambient air into the chamber.

### Exhaust

The exhaust port is typically located below the chamber’s control panel, where process air can be expelled without blowing directly onto personnel and where freshly drawn-in air cannot short-cycle directly out of the environment.

A fan or blower is placed within the exhaust port housing, and pulls process air out of the chamber when energized.  The exhaust port also includes sheet metal covers which control the direction of exhaust flow and prevent light contamination into and out of the environment.  The exhaust port can be closed with an internal slide cover to restrict the flow of ambient air into the chamber, similar to the intake.

### Exchange Rate

The fresh air exchange system is designed to provide between three (3) to eight (8) volumetric air exchanges per hour of the chamber environment, when enabled.

### Operation

The system is controlled by an on/off switch on the chamber’s control panel.  When set to the ‘on’ position, the exhaust fan is energized and runs at a constant rate.  When in the ‘off’ position, the fresh air fan is de-energized.  Please note that the system may draw some fresh air into the chamber through open intake and exhaust ports when the system is turned off.  To minimize fresh air ingress, turn off the fresh air system and close the internal slide covers on both ports.

It is recommended that the fresh air exchange system be deactivated below chamber process temperatures of 10°C.  Below this temperature, it is assumed that personnel is not occupying the chamber environment for long durations, and that respiration rates of biological specimens are reduced.

Operating the fresh air exchange system below 10°C can place a significant demand upon the temperature control system, which can shorten its lifespan and lead to temperature control problems.  Warm ambient air drawn into a cold environment can also create condensation on chamber walls and sensitive chamber components.

### Condensation

Fresh ports are generally less insulated than the chamber walls into which they are installed and are thus prone to the formation of condensation on their internal or external surfaces.  The degree to which condensation may form is dependent upon conditions inside the controlled environment relative to ambient.

## Sensor Aspirator Box

All sensors used to monitor and control environmental chamber parameters are typically housed inside an aspirator box.  This enclosure serves to protect the sensors from dirt and debris, allows for their placement in different locations throughout the environment, and insulates them from the effects of light radiation on calibration offsets.

The aspirator box assembly’s fan draws chamber process air across the sensors constantly to provide a stable sampling of environmental conditions.  Whenever the chamber’s control system is turned on, the aspirator fan is energized.

Aspirator boxes often have sufficiently long wire tethers to allow for their placement close to experiments.  The cabling is limited by the sensors and cannot reach all areas of the chamber.

Note: When placing an aspirator assembly inside a control environment, make sure flow through the aspirator is unobstructed and in the same direction as the process air flow (as previously described).  Obstructed, or incorrectly aligned air flow may cause control problems in the environment.

## Dry Contacts

A terminal block is provided on top of the control box to connect an existing remote alarm system to the chamber’s temperature safety system.  Dry alarm contacts are provided with a normally closed set of contacts and a normally open set of contacts.  Depending on the input trigger required for the existing alarm system, make an electrical connection from the alarm input to the appropriate terminals on the terminal block.  If a temperature safety occurs, the state at each set of the contacts will change, meaning that the normally open contacts will close and the normally closed contacts will open.

# Maintenance

Performing regular maintenance on an environmental chamber reduces performance degradation, improves reliability and extends operational life.  In addition to any user-specific maintenance protocols, the following general maintenance tasks should be performed at the recommended intervals.

Turn chamber power off prior to performing any maintenance inside the environment.

Disconnect chamber supply power prior to performing any maintenance inside the chamber’s mechanical compartment or control box.

Unplug any subassemblies to be serviced inside the unit prior to performing any maintenance.

## General Maintenance Summary Table

| Maintenance Task | Recommended Interval Schedule |
| --- | --- |
| Clean the Chamber | Every 3 months, or as needed |
| Clean the Condenser | Every 3 months, or as needed |
| Clean the Evaporator | Every 12 months |

## Cleaning the Chamber

Wipe down all interior and exterior surfaces with a solution of water and mild detergent.  Do not use abrasive agents or cleaning pads as these could scratch surfaces.

Glass cleaners can be used on stainless steel surfaces to prevent streaking.

## Clean the Heat Exchangers

Air must pass freely through the evaporator and condenser for proper performance.  Ambient air is drawn into and forced through the condenser by the condensing unit fan.  Chamber process air is drawn into and forced through the evaporator by a bank of fans inside the evaporator cover.  In both cases, the fans run continuously while the chamber is in operation and there are no active alarms.

Power down the chamber prior to any cleaning effort on refrigeration system components.

Please be aware that some parts of the condensing unit become hot while the unit is in operation and can remain hot for some time after the unit is powered down.

### Clean the Condenser (Air Cooled)

For air cooled refrigeration systems, regular maintenance on the condenser is required for proper operation.  The condenser is the primary heat exchanger for the refrigeration system to reject heat from the system.  If the coils become dirty or blocked with debris, the refrigeration system’s efficiency will decrease, affecting temperature control.

Once powered down, remove any dust or debris collected upon the condenser fins, copper tubing, fan blades and motor.  Blow dust out of the condenser from behind with compressed air, a brush, or vacuum and wipe dust from the fan blades.

Air cooled condensing units on walk-in chambers are located on top of the unit, adjacent to the unit, or remote from the unit – often outside.  To access the condenser coils, turn off chamber power, disconnect electrical services at the condensing unit and remove any external cover panels.

### Clean the Evaporator

Once the unit is powered down, remove the evaporator cover to access the evaporator coil and air circulation fans.

- Disconnect the flexible condensate drain line from the evaporator housing

- Remove the sheet metal screws that secure the evaporator cover to the main housing or support brackets, depending on the incubator model.

- Disconnect the electrical connection(s) between the fans affixed to the cover and main housing

A picture of a walk-in evaporator coil assembly with dehumidification is shown below for reference.

Remove screws at each end to drop pan

Disconnect drain lines

Clean evaporators and fan blades

Remove drip deflectors

Drain pan

Remove cover panels indicated in the previous diagram, then perform a visual inspection of the evaporator coil assembly:

- Look for signs of corrosion on the copper tubing, aluminum heat exchanger fins, housing and solder joints

- Look for stains or discoloration on headers, return bends, coils, and fins.  If stains or discoloration is evident, check suspect areas with a refrigerant leak detector

- Check drain pans and condensate drain lines and fittings.  Make sure these are clear of debris that could inhibit the free flow of waste condensation out of the system

- Make sure fan blades are balanced and show no signs of premature wear

- Tighten any loose fasters that affix the evaporator fans to the housing

If the evaporator has a brown phenolic coating and/or a white painted housing as shown in the photo above, please refer to the supplemental Q9 Phenolic Coating User Manual for specific cleaning requirements in lieu of these general guidelines.

Clean the evaporator coil by dislodging dirt and debris via soft brush, compressed air or vacuum, and inspect the condensate drain line for any cracks, wear or blockages.

Remove dust and debris from the evaporator fans and wipe any dirt or residue from inside the evaporator cover and housing.

Inspect all electrical wiring and components

- Perform a visual inspection for any abrasions, wear, kinks, discoloration, or exposed copper wire filament and replace any damaged areas

- Verify that all electrical ground connections are secure, and tighten where necessary

### Inspect the Aspirator Box Assembly

- Inspect the housing and remove any dust or debris which could inhibit the proper flow of air over the internal sensor suite

- Remove the cover and wipe dust and debris away from the sensors with a soft cloth

- Inspect the electrical cabling for any abrasions, wear, kinks, discoloration, or exposed copper wire filament and replace any damaged areas

- Inspect the aspirator box when the chamber is powered on to ensure the internal fan is operational

# Troubleshooting

## Loss of Temperature Control – General

| Possible Reason | Solution |
| --- | --- |
| Improper or incorrect signal between controller and control relay(s) | Check for ~5VDC output signal between the controller and solid-state relay. Output state is indicated by LEDs on the controller overlay. Contact factory if insufficient output voltage from the control system |
| Temperature sensor | Check the sensor for correct output signal. Refer to the RTD temperature vs. resistance table in the controller manual. Contact factory if sensor output does not match the table values. |
| Solid-state relays | If trigger voltage is correct to the relay coil, check for the correct output voltage through the relay. Refer to the chamber’s electrical schematic for the VAC output to the affected heating and/or cooling device(s) |
| Hot gas solenoid valve | When in heating mode, the hot gas solenoid should be in the open position allowing hot gas to flow into the evaporator coil. Check the solenoid valve for proper actuation and the valve body for internal blockages. |
| Liquid line solenoid valve | When in cooling mode, the liquid line solenoid valve should be in the open position, allowing liquid refrigerant to flow into the thermal expansion device(s). Check the solenoid for proper actuation and the valve body for internal blockages. |

## Loss of Temperature Control - Overheating

| Loss of refrigerant | Check system pressures and check for leaks with a leak detector. Evacuate the system, repair leak(s) and recharge the system. Refer to the unit’s refrigeration diagram and serial label. |
| --- | --- |
| Restriction in the refrigeration piping or componentry | If adequate pressure and no leaks, there may be a restriction in the unit’s liquid line solenoid valve, filter/drier or thermal expansion device. Check the line temperature directly before and after each device. There may be a blockage if the outlet is colder than the inlet to the solenoid and drier. |
| Insufficient heat exchange through the condenser | Check that the ambient air temperature is below 90°F. Make sure the condenser fan and coils are clean. Make sure the condenser fan is running. |
| Compressor motor fails to pump refrigerant throughout the system | Check electrical power to the motor and start components. Perform an ohm check on the motor windings. Check high-side system pressure if the unit runs |
| Insufficient heat exchange through the evaporator | Check for proper fan operation. Fans should be continuously running, drawing air into the evaporator. Make sure the evaporator coils are clean. |

## Loss of Temperature Control – Overcooling

| Electric heater malfunction | Check output voltage from the controller to the heater relay. Check output voltage from the heater relay to the heater element. Check the circuit breaker in the heater electrical circuit. |
| --- | --- |
| Malfunctioning hot gas solid-state relay (capillary tube systems) | Check output voltage from the controller to the hot gas line solenoid valve. Check voltage between the relay and solenoid valve. The solenoid valve should be open whenever the ‘heat’ LED is lit |
| Malfunctioning liquid solid-state relay (TXV systems) | Check output voltage between controller and relay. Check voltage between relay and solenoid. The solenoid should fully close when the ‘heat’ LED indicator light is energized. |
| Blocked hot gas solenoid valve (capillary tube systems) | If the hot gas solenoid valve gets blocked, then flow will divert through the capillary tube, causing the system to cool even when calling for heat. |

## General

| Chamber does not power on | Check facility electrical services and circuit breakers. Check that the chamber is properly plugged into the facility services. Check disconnect switch, power switch and circuit breakers inside the unit’s lower control box compartment. |
| --- | --- |
| Individual chamber component does not turn on | Check controller settings and verify that the device has been properly enabled. Check the control box for tripped circuit breakers. |
| Condensation | The amount of condensation that can form inside and be rejected from a controlled environment can fluctuate, and is highly dependent upon ambient conditions, interior environmental conditions, fresh air exchange through air ports, frequency of door openings, product load and optional humidity control systems. Minimize door openings and the amount of moisture present inside the unit, and close the access port(s) |

## Parts and Service

Replacement parts can be ordered from Percival Scientific at:

Ph: 1-800-695-2743

Email: parts@percival-scientific.com

For additional questions regarding this product, please contact Percival’s customer support team at:

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com