H11/H14 – Ultrasonic Humidifier

Operator’s Manual

505 Research Drive  Perry, Iowa 50220

Toll-free 800.695.2743  515.465.9363  fax 515.465.9464

www.percival-scientific.com

[This page intentionally left blank]

# Introduction

To help customers create better science, Percival Scientific continuously studies and refines the design of its products to provide the best possible solutions and equipment in the industry.  As such, changes in the design and/or operation of equipment may occur without publication.  Please feel free to contact Percival Scientific directly for product updates.

Information contained within this supplemental user manual shall supersede any conflicting information contained within the basic chamber user manual.

Please read this manual carefully before proceeding with operation

# Revision Log

| Version | Date | Description |
| --- | --- | --- |
| Rev A | 10/14/2024 | Initial Release |

Introduction	3

Revision Log	4

Warranty Information	5

Definitions	8

General	8

Performance Criteria	9

Performance Specifications	9

Process Value Stability	9

Relative Humidity Control Range (Incubators)	9

Relative Humidity Control Range (Plant Growth)	9

Relative Humidity Control Accuracy	9

Water Quality Requirements	10

Water Usage	10

Installation Guidelines	10

Additive Humidification Components Provided	11

Site Requirements	11

Ultrasonic Humidifier Installation	11

Water Supply Line Installation	16

Fill the Humidifier with Water	16

Operation	17

General	17

Additive Humidification	17

Relative Humidity Control Envelope	18

Maintenance	19

Clean the humidifier module	19

Troubleshooting	20

Parts and Service	21

# Warranty Information

Equipment Limited Warranty

Thank you for purchasing equipment from Percival Scientific, Inc.  Our goal is to provide every customer with an exceptional ownership experience.  We are committed to providing products and services of the highest quality, and your satisfaction is our top priority.  In support of this objective, we present the following warranty.

Warranty Terms

Percival Scientific, Inc. (hereinafter referred to as Percival) warrants that equipment manufactured by Percival shall be free from defect in material or workmanship, when installed, operated and maintained in accordance with instructions provided by Percival.

Warranty Coverage

- Coverage period shall be twenty-four (24) months from the date of delivery

- Percival’s obligation under this warranty is limited to the repair or replacement of defect at no charge for parts, shipping and labor

- Percival reserves the right to inspect the equipment after a warranty claim is made to determine the proper course of corrective action

Exclusions

- Parts or elements damaged from improper installation, including mechanical, electrical and plumbing services, IT infrastructure, inadequate room ventilation, improper water quality, and ambient conditions that are out of specification

- Parts or elements damaged from improper operation

- Parts or elements damaged from improper maintenance

- Parts or elements considered to be consumable, that outlive their normal expected terms of usefulness, such as: light bulbs, filters, desiccant material and CO2 scrubbing media

- Damage resulting from shipment to or relocation within the customer’s facility after equipment has been signed for and ownership is transferred to the customer

- Modification to equipment by the customer without Percival’s consent may void this warranty

- Repair work undertaken without prior consent of Percival may void this warranty

- Damage caused by accident, fire, Force Majeure or Acts of God

Additional Terms

- This warranty is in lieu of any other warranties, expressed or implied, including merchantability or fitness for a particular purpose

- Warranty terms described herein cannot be changed without the expressed, written consent from Percival

- The owner agrees that Percival’s sole liability with respect to defective parts or materials shall be set forth in this warranty, and any claims for incidental or consequential damages are expressly excluded

- This warranty is non-transferable and applies only to the original purchaser

- Repaired or replaced parts shall be warranted only for the unexpired portion of the original warranty, unless otherwise defined

- Parts replaced under warranty shall become the property of Percival

To Make a Warranty Claim

- Document the issue with detailed notes and photographs along with the equipment’s serial number

- Contact Percival’s customer support team at customerservice@percival-scientific.com, parts@percival-scientific.com or 1-800-695-2743

- Once your claim is received, our customer support team will review it and contact you with further instructions

# Definitions

‘Relative Humidity’ and ‘%RH’ are used interchangeably in this manual.  At a given temperature, and pressure, air can hold only a certain amount of water vapor.  Relative humidity expresses the amount of water vapor present in air as a percentage of the amount at full saturation.

# General

Relative humidity is an important variable in many controlled environment applications, and its precise control is crucial for maintaining the integrity and reliability of experimental conditions. Proper humidity levels are essential for simulating real-world environments, ensuring accurate data collection, and preserving the quality of sensitive materials and products.

Effective control of relative humidity not only enhances the reproducibility of experiments but also safeguards the longevity and efficacy of products, ultimately contributing to the success of research and development efforts.

Precise humidity control is vital for optimizing growth conditions, promoting healthy development, and ensuring successful germination and propagation of plants. For plant growth chamber applications, humidity regulation helps prevent dehydration and supports the physiological processes necessary for robust plant growth. Similarly, in biological incubation applications, maintaining appropriate humidity levels is essential for the growth of microorganisms, cell cultures, and other biological specimens, as it directly influences metabolic activities and overall viability.

In other applications, such as pharmaceuticals, food storage, and material testing, fluctuations in humidity can lead to degradation, contamination, or altered performance of samples.

H11 and H14 (Ultrasonic Humidifier) options provide the capability to control relative humidity process values inside Percival Scientific environmental chambers.  The system increases %RH via one or more ultrasonic humidifier modules and reduces %RH via a dedicated refrigeration-based evaporator coil.

This manual is intended to provide instructions on the installation, operation, maintenance and troubleshooting of H11 and H14 options for Percival Scientific’s control environmental chambers.

# Performance Criteria

Performance of the H11 and H14 (Ultrasonic Humidifier) humidity control system shall be based upon the following criteria:

- Ambient conditions surrounding the controlled environment of 24°C (75°F) and 55%RH

- Empty chamber (i.e. no product load)

- A single point of control in the location of the %RH sensor (typically located near the evaporator intake, or at shelf level inside an aspirated sensor box)

- An understanding that door openings, changes to the chamber’s control system or physical settings can temporarily destabilize temperature and %RH levels, causing process values to deviate from their specified tolerance range for control

- Relative humidity levels inside the controlled environment shall be limited by the dew points of the primary (temperature) and secondary (dehumidification) evaporator coils

# Performance Specifications

## Process Value Stability

Process value stability shall be achieved after 30 minutes within reaching the temperature and relative humidity set points.  After achieving stable operating conditions within the environment, 95% (2σ) of the %RH process values, measured under the conditions below, shall fall within the specified tolerance ranges below.

## Relative Humidity Control Range (Incubators)

Above ambient to 90%RH, between 15 – 30°C (lights off)

Above ambient to 80%RH, between 15 – 30°C (lights on)

## Relative Humidity Control Range (Plant Growth)

Above ambient to 90%RH, between 15 – 30°C (lights off)

Above ambient to 70%RH, between 15 – 30°C (lights on)

## Relative Humidity Control Accuracy

H11 System: ± 5%RH, around the humidity set point

H14 System: ± 10%RH, around the humidity set point

## Water Quality Requirements

pH: 7.0 ± 0.5

Filtration: <2 microns

Resistivity: 0.5 – 5.0 MΩ (<20 µS)

Supply Pressure: 7 – 87 psi (0.5 – 6.0 bar)

Supply Temperature: 41 – 104°C (5 – 40°C)

## Water Usage

Estimated water usage varies by ultrasonic humidifier type and the quantity used.  Please refer to the original project specifications or quotation for estimated humidifier water consumption rates.

The following table includes per-device water usage for the ultrasonic humidifier models commonly included with H11 and H14 optional equipment packages:

| Manufacturer | Model | lbs/hr | L/hr |
| --- | --- | --- | --- |
| Stulz | FN-400H | 1.1 | 0.5 |
| Stulz | SCA-1000 | 2.2 | 1.0 |
| Percival | EHU-130 | 3.3 | 1.5 |
| Stulz | SCA-2000 | 4.4 | 2.0 |

# Installation Guidelines

H11 and H14 optional equipment packages are provided with all hardware, electrical infrastructure and control system configurations required for integration with most controlled environment chambers from Percival.  Please note that the specific implementation of these options will differ according to the base chamber model into which the humidity system is added.  Despite these differences, the general concept, function and control of the H11 or H14 system remains the same across all variations.

Proceed with all general chamber installation guidelines contained within the base chamber’s user manual before proceeding with the following humidifier installation guidelines.

## Additive Humidification Components Provided

- Vaisala HMP60 series (included with H14) or HMP110 series (included with H11) relative humidity sensor

- Ultrasonic humidifier module(s)

- Flexible tubing

- Drip and overflow collection pan(s) with drain(s)

- Solid-state relay(s), to facilitate control of humidification system element(s)

- Power supply or power transformer

- Resettable branch circuit protection

- Water supply shutoff valve

- Dedicated electrical connections through which the humidifier(s) connect to the chamber’s control system

- Two (2) additional access ports with plugs, for fresh air exchange

  - One (1) port located on the lower, left side of the chamber

  - One (1) port located on the rear wall, behind the evaporator coil

## Site Requirements

The end user is responsible for providing a water source per the water volume, quality and supply pressure requirements outlined in the previous section of this manual.  The water source should have sufficient capacity to initially fill the humidifier(s), then maintain the appropriate water level while the chamber and humidification system are in operation.

## Ultrasonic Humidifier Installation

Electrical connection points for humidifiers are expressly labelled as such.  Do not connect humidifier plugs to any differently marked, or unmarked receptacles.  Failure to follow these instructions may cause malfunction and/or damage to the humidifier assembly or chamber.  Please contact Percival Scientific with any installation questions.

When possible, ultrasonic humidifier assemblies ship pre-installed on Percival chambers and require no installation. However, some ultrasonic humidifier applications require components that cannot fit within a chamber’s mechanical compartment and must be installed externally.

Externally mounted ultrasonic humidifier assemblies large enough to prevent a chamber from fitting through a standard doorway are disconnected after validation testing, packaged, and shipped inside the unit.

Any ultrasonic humidifiers shipped inside a chamber as uninstalled parts require installation prior to first use. To install:

- Remove the humidifier assembly and associated parts from their packaging.

- Determine the location at which the humidifier(s) are to be installed and the proper orientation.  The table below describes the general size and location for ultrasonic humidifiers for each Percival chamber model, and some reference pictures:

| Model | Humidifier Type | Location |
| --- | --- | --- |
| 22 Series | FN-400H | External, RHS |
| 30 Series | FN-400H | External, rear or top |
| 36 Series | FN-400H | External, top |
| 41 Series | SCA-1000 | External, top |
| 66 Series | SCA-1000 | Internal, top |
| 75 Series | SCA-1000 or SCA-2000 | External, top |
| 95 Series | SCA-2000 | External, top |
| 100 Series | SCA-2000 | Internal, top |
| PGC Series | SCA-1000 or SCA-2000 | Internal, side mechanical compartment or below perforated floor |
| Walk-ins | SCA-2000 | Consult Factory |

The following images are provided for general reference regarding various ultrasonic humidifier mounting locations across different Percival product lines:

(FN-400H, top-mounted on a 36 Series chamber)

(FN-400H, rear-mounted on a 30 Series chamber)

(FN-400H, side-mounted on a 22 Series chamber)

(SCA-1000, top-mounted)

(SCA-2000, top-mounted)

- If included in the ultrasonic humidifier parts kit, install the drip pan and/or ultrasonic humidifier mounting bracket(s)

- Place the ultrasonic humidifier assembly in its proper location and secure it in place with provided mounting brackets, hardware and fasteners.

- Connect a ¼” o.d. water supply line to the provided shutoff valve (shown below), then to each humidifier’s water inlet fitting.  Ensure that the water line is fully inserted into any push-in fittings and properly tightened into any compression fittings.

(1/4-turn water shutoff valve, shown installed in-line, in the ‘open’ position)

- Connect tubing between the humidifier inlet/outlet(s) and chamber outlet/inlet connection point(s).  If chamber connection points are located inside the mechanical compartment, tubing is typically left connected to the chamber and coiled inside the mechanical compartment for shipping.  Ensure tubing is tightly secured to prevent water leaks using provided hose clamps.

- Plug the humidifier’s power cord(s) into the dedicated receptacle(s) or connectors.  All dedicated receptacles for the humidifier will be labelled ‘Humidifier’ to indicate their intended function.  Make sure all plugs are fully seated into their corresponding receptacles or connectors.

## Water Supply Line Installation

It is highly recommended that the water supply line be purged of any debris prior to connection with the chamber’s humidity system.  Even very small pieces of debris in the supply line can get lodged inside the fill valve, preventing it from properly closing.

All types of ultrasonic humidifiers require a water supply for proper operation.  Do not operate the chamber’s humidity system unless supply water has been properly connected and the humidifier reservoirs are filled with enough water to cover the water atomizing transducers.

To make the supply water connection:

- Connect supply water line to the provided, quarter-turn shutoff valve.  In some cases, this valve is affixed to the chamber, and in other cases, it’s provided as a loose item.

- The shutoff valve requires a ¼” o.d. supply line.  To connect this line to the valve, insert it into the valve body until it hits a positive stop.  The water valve’s push-type connection fitting will provide some resistance to the tubing before it properly seats.  Approximately a 5/8” length of the supply line tubing should extend into the valve when properly inserted.

## Fill the Humidifier with Water

When power and water are provided, all available ultrasonic humidifier modules automatically fill with water and maintain the appropriate water level for proper operation.  Ultrasonic transducers will not energize unless immersed in the correct amount of water.

Please refer to the ultrasonic humidifier manufacturer’s supplemental user manual for additional information.

# Operation

## General

Relative humidity set points are achieved through a closed-loop control scenario.  The chamber’s control system continuously measures the relative humidity process value inside the controlled environment.  This process value input from the sensor is compared to the prescribed set point.  The control system adjusts operation of the humidifier, depending on the difference between the values.

Additive humidification is accomplished with an ultrasonic humidifier.  The humidifier has a water reservoir which maintains a specific water level.  Transducers located inside the reservoir vibrate at a specific frequency and amplitude to atomize water into very small droplets.

If the process value is lower than the set point, the humidifier assembly’s atomizing transducers are energized.

If the process value is higher than the set point, the humidifier assembly’s atomizing transducers are de-energized

The on/off ratio of the %RH control system outputs are determined by an independent PID control loop, which helps balance system responsiveness with control accuracy.  For example, if the process value is significantly lower than the set point, the control system will energize the humidifier for a large percentage of the cycle time.  As the process value increases, approaching the set point, the system will reduce the amount of time the humidifier is energized to prevent overshoot.

## Additive Humidification

H11 and H14 (Ultrasonic Humidifier) systems raise the %RH level inside Percival controlled environmental chambers by injecting atomized water particles into the chamber’s process air stream.

To operate the humidity control system:

- Ensure the humidifier is installed according to the previous installation steps

- Open the humidifier water supply valve and turn any power switched on the humidifier module to the ‘on’ or ‘|’ position

- Enable Channel A of the %RH control loop in the Inputs Menu of the control system.

Please refer to the control system user manual for specific instruction on this process

- Set the desired relative humidity level via a Manual Set Point, or Multi-step Program.  Please refer to the control system user manual for specific instructions on this process

The %RH level inside the environment will self-adjust to the desired set point, provided the operational conditions described earlier in this manual are all satisfied.

Due to the psychrometric conditions within the controlled environment, achievement of the set point may take multiple hours.

## Relative Humidity Control Envelope

The chart below describes the approximate performance envelope for relative humidity control in most Percival controlled environment chambers.

H11 and H14 humidity control options are additive only.  Only the upper %RH boundary of the performance envelope can be achieved with these options.

As described in the chart above, it is possible to prescribe temperature and relative humidity parameter pairs that fall outside of the performance specification envelope.  Within the orange region of the performance envelope, achievement of the set point and control stability cannot be guaranteed.  Within this region, there also exists an increased likelihood of excessive condensation, loss of process control, reduced part lifespan, and chamber malfunction.

Please note that the lower %RH boundaries of the chart above are only achievable with optional dehumidification systems.

As a protective measure, the chamber’s control system automatically disables the humidification system if the set point is less than or equal to 10°C or greater than or equal to 40°C.

# Maintenance

## Clean the humidifier module

- Shut off supply water to the chamber

- If present, turn off the humidifier’s power switch

- Unplug the humidifier’s electrical cord

- Disconnect the humidifier’s exhaust and intake ductwork

- Remove humidifier cover(s)

- Remove the humidifier from its support bracket(s) and empty water from the reservoir

- Remove any debris that may have fallen or migrated into the reservoir

- Clean the humidifier as directed in the unit’s user manual

- Inspect the atomizing element(s) for corrosion, cracks or discoloration

- Check for loose or damaged wiring during the cleaning process

- Re-install the humidifier components, cover(s), ductwork and connect the water feed line and power cord

- Open the water supply valve

Recommended Cleaning Schedule: as needed between experiments, or as prescribed in the humidifier module operating instructions

Recommendations for Storage: Completely drain water from the humidifier prior to any periods where the chamber will not be in use.

# Troubleshooting

| Problem | Reason | Solution |
| --- | --- | --- |
| Reservoir won’t fill | Interrupted water supply | Check supply line for blockages and valve position |
|  | Non-opening fill valve | Clear blockage at valve inlet Check level sensor and solenoid valve electrical system |
| Low reservoir water level | Level sensor is out of adjustment | Check position of the level sensor and adjust as needed Check inlet valve for blockage, preventing it from closing |
| High reservoir water level |  |  |
| Atomizers are not energizing | Electrical issue | Check electrical connections Check the humidity system’s circuit breaker in the control box |
|  | Water level is too low inside humidifier reservoir | Check level sensor and water filling system |
| Mineral deposits on chamber surfaces | Water quality | Check supply water quality against system requirements, specified above |
| The relative humidity level won’t increase or decrease inside the chamber | Control system setup | Make sure Channels A and B are enabled in the RH inputs menu Set points may be outside of allowable operating range |
|  | Electrical | Check for tripped circuit breaker(s) in the control panel Make sure power cords are connected to appropriate receptacle and are fully seated |
|  | Mechanical | Check humidifier water level Check water supply and shutoff valve position |
| Water is collecting outside of the ultrasonic humidifier | Mechanical | Check for tightness of humidifier lid, humidity injection tubing and water supply line connections |

# Parts and Service

Replacement port plugs and caps can be ordered from Percival Scientific at:

Ph: 1-800-695-2743

Email: parts@percival-scientific.com

For additional questions regarding this product, please contact Percival’s customer support team at:

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com