H17 – Dehumidifier Only

Operator’s Manual

505 Research Drive  Perry, Iowa 50220

Toll-free 800.695.2743  515.465.9363  fax 515.465.9464

www.percival-scientific.com

[This page intentionally left blank]

# Introduction

To help customers create better science, Percival Scientific continuously studies and refines the design of its products to provide the best possible solutions and equipment in the industry.  As such, changes in the design and/or operation of equipment may occur without publication.  Please feel free to contact Percival Scientific directly for product updates.

Information contained within this supplemental user manual shall supersede any conflicting information contained within the basic chamber user manual.

Please read this manual carefully before proceeding with operation

# Revision Log

| Version | Date | Description |
| --- | --- | --- |
| Rev A | 6/19/2025 | Initial Release |

Introduction	3

Revision Log	4

Warranty Information	6

Definitions	8

General	8

Performance Criteria	9

Performance Specifications	9

Process Value Stability	9

Relative Humidity Control Range (Incubators and Plant Growth, except 22 and 30 Series)	9

Relative Humidity Control Accuracy	9

Installation Guidelines	10

Dehumidification Components Provided	10

Site Requirements	10

Installation of Dehumidification Components	11

Operation	11

General	11

Dehumidification	12

Relative Humidity Control Envelope	13

Maintenance	14

Clean the Dehumidification Evaporator Coils	14

Troubleshooting	15

Parts and Service	16

# Warranty Information

Equipment Limited Warranty

Thank you for purchasing equipment from Percival Scientific, Inc.  Our goal is to provide every customer with an exceptional ownership experience.  We are committed to providing products and services of the highest quality, and your satisfaction is our top priority.  In support of this objective, we present the following warranty.

Warranty Terms

Percival Scientific, Inc. (hereinafter referred to as Percival) warrants that equipment manufactured by Percival shall be free from defect in material or workmanship, when installed, operated and maintained in accordance with instructions provided by Percival.

Warranty Coverage

- Coverage period shall be twenty-four (24) months from the date of delivery

- Percival’s obligation under this warranty is limited to the repair or replacement of defect at no charge for parts, shipping and labor

- Percival reserves the right to inspect the equipment after a warranty claim is made to determine the proper course of corrective action

Exclusions

- Parts or elements damaged from improper installation, including mechanical, electrical and plumbing services, IT infrastructure, inadequate room ventilation, improper water quality, and ambient conditions that are out of specification

- Parts or elements damaged from improper operation

- Parts or elements damaged from improper maintenance

- Parts or elements considered to be consumable, that outlive their normal expected terms of usefulness, such as: light bulbs, filters, desiccant material and CO2 scrubbing media

- Damage resulting from shipment to or relocation within the customer’s facility after equipment has been signed for and ownership is transferred to the customer

- Modification to equipment by the customer without Percival’s consent may void this warranty

- Repair work undertaken without prior consent of Percival may void this warranty

- Damage caused by accident, fire, Force Majeure or Acts of God

Additional Terms

- This warranty is in lieu of any other warranties, expressed or implied, including merchantability or fitness for a particular purpose

- Warranty terms described herein cannot be changed without the expressed, written consent from Percival

- The owner agrees that Percival’s sole liability with respect to defective parts or materials shall be set forth in this warranty, and any claims for incidental or consequential damages are expressly excluded

- This warranty is non-transferable and applies only to the original purchaser

- Repaired or replaced parts shall be warranted only for the unexpired portion of the original warranty, unless otherwise defined

- Parts replaced under warranty shall become the property of Percival

To Make a Warranty Claim

- Document the issue with detailed notes and photographs along with the equipment’s serial number

- Contact Percival’s customer support team at customerservice@percival-scientific.com, parts@percival-scientific.com or 1-800-695-2743

- Once your claim is received, our customer support team will review it and contact you with further instructions

# Definitions

‘Relative Humidity’ and ‘%RH’ are used interchangeably in this manual.  At a given temperature, and pressure, air can hold only a certain amount of water vapor.

Relative humidity expresses the amount of water vapor present in air as a percentage of the amount at full saturation.

# General

Relative humidity is an important variable in many controlled environment applications, and its precise control is crucial for maintaining the integrity and reliability of experimental conditions. Proper humidity levels are essential for simulating real-world environments, ensuring accurate data collection, and preserving the quality of sensitive materials and products.

Effective control of relative humidity not only enhances the reproducibility of experiments but also safeguards the longevity and efficacy of products, ultimately contributing to the success of research and development efforts.

Precise humidity control is vital for optimizing growth conditions, promoting healthy development, and ensuring successful germination and propagation of plants. For plant growth chamber applications, humidity regulation helps prevent dehydration and supports the physiological processes necessary for robust plant growth. Similarly, in biological incubation applications, maintaining appropriate humidity levels is essential for the growth of microorganisms, cell cultures, and other biological specimens, as it directly influences metabolic activities and overall viability.

In other applications, such as pharmaceuticals, food storage, and material testing, fluctuations in humidity can lead to degradation, contamination, or altered performance of samples.

The H17 (Dehumidifier Only) option provides the capability to control relative humidity process values inside Percival Scientific environmental chambers.  The system reduces %RH via a dedicated refrigeration-based evaporator coil, but has not ability to actively increase %RH levels.

This manual is intended to provide instructions on the installation, operation, maintenance and troubleshooting of the H17 (Dehumidifier Only) option for Percival Scientific’s control environmental chambers.

# Performance Criteria

Performance of the H17 (Dehumidifier Only) humidity control system shall be based upon the following criteria:

- Ambient conditions surrounding the controlled environment of 24°C (75°F) and 55%RH

- Empty chamber (i.e. no product load)

- A single point of control in the location of the %RH sensor (typically located near the evaporator intake, or at shelf level inside an aspirated sensor box)

- An understanding that door openings, changes to the chamber’s control system or physical settings can temporarily destabilize temperature and %RH levels, causing process values to deviate from their specified tolerance range for control

- Relative humidity levels inside the controlled environment shall be limited by the dew points of the primary (temperature) and secondary (dehumidification) evaporator coils

# Performance Specifications

## Process Value Stability

Process value stability shall be achieved after 30 minutes within reaching the temperature and relative humidity set points.  After achieving stable operating conditions within the environment, 95% (2σ) of the %RH process values, measured under the conditions below, shall fall within the specified tolerance ranges below.

## Relative Humidity Control Range (Incubators and Plant Growth, except 22 and 30 Series)

Down to 40%RH, between 15 – 30°C (lights off)

Down to 40%RH, between 15 – 30°C (lights on)

22 and 30 Series chambers of all types are limited by the physical size of their refrigeration systems and cannot achieve lower than 50%RH, between 15-30°C (lights on and off) at the ambient operating conditions specified above

## Relative Humidity Control Accuracy

H17 System: ± 10%RH, around the humidity set point

# Installation Guidelines

The H17 optional equipment package is provided with all hardware, electrical infrastructure and control system configurations required for integration with most controlled environment chambers from Percival.  Please note that the specific implementation of this option will differ according to the base chamber model into which the humidity system is added.  Despite these differences, the general concept, function and control of the H17 system remains the same across all variations.

Proceed with all general chamber installation guidelines contained within the base chamber’s user manual before proceeding with the following humidifier installation guidelines.

## Dehumidification Components Provided

- One or more bare-tube dehumidification evaporator coils with a thermal expansion device such as a capillary tube or expansion valve

- A solenoid valve to control refrigerant flow to the dehumidification evaporator(s), through their thermal expansion device(s)

- Electric heating element(s), for heating air, but suitable for wet environments

- Solid-state relay(s), to facilitate control of dehumidification refrigeration solenoid valve and electric air heater(s)

- Resettable branch circuit protection

- A means to capture and drain away waste condensate from the dehumidification evaporator(s)

## Site Requirements

The end user is responsible for providing a means by which condensate rejected by the controlled environment can be managed.  Condensate can be captured in a reservoir and drained manually, or routed to a facility drain directly or via a condensate pump.

Please contact Percival Scientific with any questions or for the purchase of optional drain system components.

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com

## Installation of Dehumidification Components

All provided dehumidification components described above are pre-installed and fully integrated into every chamber model at Percival’s factory prior to shipment.  No installation is required by the end user.

# Operation

## General

Relative humidity set points are achieved through a closed-loop control scenario.  The chamber’s control system continuously measures the relative humidity process value inside the controlled environment.  This process value input from the sensor is compared to the prescribed set point.  The control system adjusts operation of the dehumidifier, depending on the difference between the values.

If the %RH process value is lower than the set point:

- The dehumidification system’s refrigeration solenoid valve(s) and electric air heater(s) are de-energized

- The %RH level inside the controlled environment will passively increase over time until equilibrium is reached.  Relative humidity levels above the equilibrium point - which is dependent upon ambient conditions, fresh air exchange and product load characteristics - cannot be achieved without an additive humidification system.

If the %RH process value is higher than the set point:

- The dehumidification system’s refrigeration solenoid valve(s) and electric air heater(s) are energized

- The system will lower the %RH level inside the unit set point or system performance specification limit(s) have been reached

The on/off ratio of the %RH control system outputs are determined by independent PID control loops, which help balance system responsiveness with control accuracy.  For example, if the process value is significantly higher than the set point, the control system will energize the dehumidifier for a large percentage of the cycle time.  As the process value decreases, approaching the set point, the system will begin to reduce the amount of time the dehumidifier is energized.  Because H17 does not include an additive %RH system, there exists a greater potential for the %RH process value to overshoot the set point and require additional time to recover.

## Dehumidification

The H17 dehumidification system reduces relative humidity inside a chamber environment by condensing moisture out of the air onto cold, bare-tube evaporator coils.

Process air is passed through the main evaporator(s) for temperature conditioning, then through the dehumidification evaporator coil(s).  When enabled, the dehumidification evaporator operates at just above the freezing point.  As moisture condenses upon the cold coil surface, it sloughs off into a drip pan and is routed to the chamber floor drain.

Upon exiting the dehumidification evaporator(s), process air is passed over electric air heaters, where heat lost from the dehumidification process is reintroduced.

When the control system calls for heating and dehumidification, refrigerant only flows through the dehumidification evaporator(s).  Heating is accomplished solely through the chamber’s main electric air heaters during this condition.

To operate the humidity control system:

- Ensure the chamber is installed according to the previous installation steps

- Enable Channels B of the %RH control loop in the Inputs Menu of the control system.

Please refer to the control system user manual for specific instruction on this process

- Set the desired relative humidity level via a Manual Set Point, or Multi-step Program.  Please refer to the control system user manual for specific instructions on this process

The %RH level inside the environment will self-adjust to the desired set point, provided the operational conditions described earlier in this manual are all satisfied.

Due to the psychrometric conditions within the controlled environment, achievement of the set point may take multiple hours.

## Relative Humidity Control Envelope

The chart below describes the approximate performance envelope for relative humidity control in most Percival controlled environment chambers.  (Per the performance specification section above, 22 and 30 series chambers are capable of dehumidification down to 50%RH).

As described in the chart above, it is possible to prescribe temperature and relative humidity parameter pairs that fall outside of the performance specification envelope.  Within the orange region of the performance envelope, achievement of the set point and control stability cannot be guaranteed.  There also exists an increased likelihood of excessive condensation, loss of process control, reduced part lifespan, and chamber malfunction within this region.

Please note that the lower %RH boundaries of the chart above are only achievable with optional dehumidification systems, and the upper boundaries are only achievable with optional additive humidification systems.

As a protective measure, the chamber’s control system automatically disables the humidification system if the set point is less than or equal to 10°C or greater than or equal to 40°C.

# Maintenance

## Clean the Dehumidification Evaporator Coils

Dehumidification evaporator coil(s) are integrated into the main evaporator assembly and constructed of the same materials.  Inspect and clean the dehumidification coils in the same manner and schedule as the main evaporator, as described in the base chamber user manual.

# Troubleshooting

| Problem | Reason | Solution |
| --- | --- | --- |
| The relative humidity level won’t decrease inside the chamber | Control system setup | Make sure Channels B is enabled in the RH inputs menu Set points may be outside of allowable operating range |
|  | Electrical | Check for tripped circuit breaker(s) in the control panel Make sure power cords are connected to an appropriate receptacle and are fully seated |
| Water is collecting outside of the chamber | Mechanical | Check drain line connection points, check for drain line termination over a facility drain. If provided, check water level on condensate pan or collection reservoir If provided, check physical drain connection to condensate pump. |
|  | Electrical | If provided, check electrical connection, power circuit and circuit breaker(s) for condensate pump |

# Parts and Service

Replacement port plugs and caps can be ordered from Percival Scientific at:

Ph: 1-800-695-2743

Email: parts@percival-scientific.com

For additional questions regarding this product, please contact Percival’s customer support team at:

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com