H3/H6 – Water Pan Humidifier with Dehumidification

Operator’s Manual

505 Research Drive  Perry, Iowa 50220

Toll-free 800.695.2743  515.465.9363  fax 515.465.9464

www.percival-scientific.com

[This page intentionally left blank]

# Introduction

To help customers create better science, Percival Scientific continuously studies and refines the design of its products to provide the best possible solutions and equipment in the industry.  As such, changes in the design and/or operation of equipment may occur without publication.  Please feel free to contact Percival Scientific directly for product updates.

Information contained within this supplemental user manual shall supersede any conflicting information contained within the basic chamber user manual.

Please read this manual carefully before proceeding with operation

# Revision Log

| Version | Date | Description |
| --- | --- | --- |
| Rev A | 10/14/2024 | Initial Release |

Introduction	3

Revision Log	4

Warranty Information	7

Definitions	9

General	9

Performance Criteria	10

Performance Specifications	10

Process Value Stability	10

Relative Humidity Control Range (Incubators, except 22 and 30 Series)	10

Relative Humidity Control Range (Plant Growth, except 22 and 30 Series)	10

Relative Humidity Control Accuracy	11

Water Quality Requirements	11

Water Usage	11

Installation Guidelines	11

Additive Humidification Components Provided	11

Dehumidification Components Provided	12

Site Requirements	12

Installation of Evaporator-mounted Reservoirs	12

Installation of Floor-mounted Reservoirs	13

Water Supply Line Installation	13

Fill the Humidifier with Water	14

Float Valve Adjustment	15

Installation of Dehumidification Components	15

Operation	15

General	15

Additive Humidification	16

Dehumidification	16

Relative Humidity Control Envelope	17

Maintenance	18

Clean the Floor-mounted System	18

Flush Evaporator-mounted System	18

Clean the Dehumidification Evaporator Coils	19

Troubleshooting	20

Parts and Service	21

# Warranty Information

Equipment Limited Warranty

Thank you for purchasing equipment from Percival Scientific, Inc.  Our goal is to provide every customer with an exceptional ownership experience.  We are committed to providing products and services of the highest quality, and your satisfaction is our top priority.  In support of this objective, we present the following warranty.

Warranty Terms

Percival Scientific, Inc. (hereinafter referred to as Percival) warrants that equipment manufactured by Percival shall be free from defect in material or workmanship, when installed, operated and maintained in accordance with instructions provided by Percival.

Warranty Coverage

- Coverage period shall be twenty-four (24) months from the date of delivery

- Percival’s obligation under this warranty is limited to the repair or replacement of defect at no charge for parts, shipping and labor

- Percival reserves the right to inspect the equipment after a warranty claim is made to determine the proper course of corrective action

Exclusions

- Parts or elements damaged from improper installation, including mechanical, electrical and plumbing services, IT infrastructure, inadequate room ventilation, improper water quality, and ambient conditions that are out of specification

- Parts or elements damaged from improper operation

- Parts or elements damaged from improper maintenance

- Parts or elements considered to be consumable, that outlive their normal expected terms of usefulness, such as: light bulbs, filters, desiccant material and CO2 scrubbing media

- Damage resulting from shipment to or relocation within the customer’s facility after equipment has been signed for and ownership is transferred to the customer

- Modification to equipment by the customer without Percival’s consent may void this warranty

- Repair work undertaken without prior consent of Percival may void this warranty

- Damage caused by accident, fire, Force Majeure or Acts of God

Additional Terms

- This warranty is in lieu of any other warranties, expressed or implied, including merchantability or fitness for a particular purpose

- Warranty terms described herein cannot be changed without the expressed, written consent from Percival

- The owner agrees that Percival’s sole liability with respect to defective parts or materials shall be set forth in this warranty, and any claims for incidental or consequential damages are expressly excluded

- This warranty is non-transferable and applies only to the original purchaser

- Repaired or replaced parts shall be warranted only for the unexpired portion of the original warranty, unless otherwise defined

- Parts replaced under warranty shall become the property of Percival

To Make a Warranty Claim

- Document the issue with detailed notes and photographs along with the equipment’s serial number

- Contact Percival’s customer support team at customerservice@percival-scientific.com, parts@percival-scientific.com or 1-800-695-2743

- Once your claim is received, our customer support team will review it and contact you with further instructions

# Definitions

‘Relative Humidity’ and ‘%RH’ are used interchangeably in this manual.  At a given temperature, and pressure, air can hold only a certain amount of water vapor.  Relative humidity expresses the amount of water vapor present in air as a percentage of the amount at full saturation.

# General

Relative humidity is an important variable in many controlled environment applications, and its precise control is crucial for maintaining the integrity and reliability of experimental conditions. Proper humidity levels are essential for simulating real-world environments, ensuring accurate data collection, and preserving the quality of sensitive materials and products.

Effective control of relative humidity not only enhances the reproducibility of experiments but also safeguards the longevity and efficacy of products, ultimately contributing to the success of research and development efforts.

In plant growth and biological incubation applications, precise humidity control is vital for optimizing growth conditions, promoting healthy development, and ensuring successful germination and propagation of plants. For plant growth chamber applications, humidity regulation helps prevent dehydration and supports the physiological processes necessary for robust plant growth. Similarly, in biological incubation applications, maintaining appropriate humidity levels is essential for the growth of microorganisms, cell cultures, and other biological specimens, as it directly influences metabolic activities and overall viability.

In other applications, such as pharmaceuticals, food storage, and material testing, fluctuations in humidity can lead to degradation, contamination, or altered performance of samples.

The H3/H6 (Water Pan Humidifier with Dehumidification) option provides the capability to control relative humidity process values inside Percival Scientific environmental chambers.  The system increases %RH via a water pan humidifier and reduces %RH via a dedicated refrigeration-based evaporator coil.

This manual is intended to provide instructions on the installation, operation, maintenance and troubleshooting of the H3/H6 (Water Pan Humidifier with Dehumidification) option for Percival Scientific’s control environmental chambers.

# Performance Criteria

Performance of the H3/H6 (Water Pan Humidifier with Dehumidification) humidity control system shall be based upon the following criteria:

- Ambient conditions surrounding the controlled environment of 24°C (75°F) and 55%RH

- Empty chamber (i.e. no product load)

- A single point of control in the location of the %RH sensor (typically located near the evaporator intake, or at shelf level inside an aspirated sensor box)

- An understanding that door openings, changes to the chamber’s control system or physical settings can temporarily destabilize temperature and %RH levels, causing process values to deviate from their specified tolerance range for control

- Relative humidity levels inside the controlled environment shall be limited by the dew points of the primary (temperature) and secondary (dehumidification) evaporator coils

# Performance Specifications

## Process Value Stability

Process value stability shall be achieved after 30 minutes within reaching the temperature and relative humidity set points.  After achieving stable operating conditions within the environment, 95% (2σ) of the %RH process values, measured under the conditions below, shall fall within the specified tolerance ranges below.

## Relative Humidity Control Range (Incubators, except 22 and 30 Series)

40 to 80%RH, between 15 – 30°C (lights off)

40 to 70%RH, between 15 – 30°C (lights on)

## Relative Humidity Control Range (Plant Growth, except 22 and 30 Series)

40 to 80%RH, between 15 – 30°C (lights off)

40 to 60%RH, between 15 – 30°C (lights on)

22 and 30 Series chambers of all types are limited by the physical size of their refrigeration systems and cannot achieve lower than 50%RH, between 15-30°C (lights on and off) at the ambient operating conditions specified above

## Relative Humidity Control Accuracy

H3 System: ± 10%RH, around the humidity set point

H6 System: ± 5%RH, around the humidity set point

## Water Quality Requirements

pH: 7.0 ± 0.5

Filtration: <2 microns

Resistivity: 0.5 – 5.0 MΩ

## Water Usage

Estimated at up to 1 gal/day

# Installation Guidelines

The H3/H6 (Water Pan Humidifier with Dehumidification) kit is provided with all hardware, electrical infrastructure and control system configurations required for integration with most controlled environment chambers from Percival.  Please note that the specific implementation of this option will differ according to the base chamber model into which the humidity system is added.  Despite these differences, the general concept, function and control of the H3/H6 system remains the same across all variations.

Proceed with all general chamber installation guidelines contained within the base chamber’s user manual before proceeding with the following humidifier installation guidelines.

## Additive Humidification Components Provided

- Vaisala HMP60 series (included with H3) or HMP110 series (included with H6) relative humidity sensor

- Water reservoir

  - Incubators utilize a floor-mounted design

  - Multi-tier chambers with solid shelving or light fixtures utilize a humidifier reservoir, mounted adjacent to (beneath) the main evaporator assembly

- Electric heating element(s), suitable for submersion in water

- Solid-state relay(s), to facilitate control of the water heating element(s)

- Resettable branch circuit protection

- Water supply shutoff valve

- A means by which a minimum water level is maintained in the water reservoir

  - Floor-mounted humidifiers utilize a mechanical float valve

  - Evaporator-mounted humidifiers utilize a supply solenoid, controlled by a level switch

- Dedicated receptacle(s) through which the humidifier reservoir’s electrical elements connect to the chamber’s control system

## Dehumidification Components Provided

- One or more bare-tube dehumidification evaporator coils with a thermal expansion device such as a capillary tube or expansion valve

- A solenoid valve to control refrigerant flow to the dehumidification evaporator(s), through their thermal expansion device(s)

- Electric heating element(s), for heating air, but suitable for wet environments

- Solid-state relay(s), to facilitate control of dehumidification refrigeration solenoid valve and electric air heater(s)

- Resettable branch circuit protection

- A means to capture and drain away waste condensate from the dehumidification evaporator(s)

## Site Requirements

The end user is responsible for providing a water source per the water quality requirements outlined in the previous section of this manual.  The water source should have sufficient capacity to initially fill the humidifier reservoir, then maintain the appropriate water level while the chamber and humidifier are in operation.

If the installation site does not have appropriate provisions for supply water, the H3/H6 system can operate in conjunction with a gravity-feed H24 (Carboy Water Reservoir) kit.  Please contact Percival Scientific for more information on this optional accessory.

## Installation of Evaporator-mounted Reservoirs

Evaporator-mounted humidifiers ship completely installed inside Percival environmental chambers.  All electrical and mechanical connections and interfaces are pre-set at Percival’s manufacturing facility, prior to shipment.

## Installation of Floor-mounted Reservoirs

Floor-mounted humidifiers are shipped inside the chamber and require installation prior to first use. To install:

- Remove the humidifier assembly from its packaging.

- Check to ensure the chamber’s condensate drain line is routed toward and inserted into the floor drain.

- Place the humidifier assembly onto the floor of the chamber, straddling the condensate drain line.

- The assembly should be positioned such that the power cord(s) are located toward the back of the chamber and the water line connection is at the front, right side of the chamber.

- Connect the ¼” o.d. water supply line to the water pan humidifier’s float valve assembly, ensuring that the water line is fully inserted and properly tightened into position.

- Plug the humidifier’s power cord(s) into the dedicated receptacle(s), found at the bottom of the wire raceway(s).  All dedicated receptacles for the humidifier will be labelled ‘Humidifier’ to indicate their intended function.  Make sure all plugs are fully seated into their corresponding receptacles.

Electrical connection points for humidifiers are expressly labelled as such.  Do not connect humidifier plugs to any differently marked, or unmarked receptacles.  Failure to follow these instructions may cause malfunction and/or damage to the humidifier assembly or chamber.  Please contact Percival Scientific with any installation questions.

## Water Supply Line Installation

It is highly recommended that the water supply line be purged of any debris prior to its connection to the chamber’s humidity system.  Even very small pieces of debris in the supply line can get lodged inside the fill valve, preventing it from properly closing.

Both evaporator and floor-mounted humidifiers require a water supply for proper operation.  Do not operate the chamber’s humidity system unless supply water has been properly connected and the humidifier pan is filled with enough water to cover the heater element(s).

To make the supply water connection:

- Locate the water supply line, located externally on the chamber.  Typically, this supply line is located on the right-hand side of the unit.

- Connect supply water to the provided, quarter-turn shutoff valve.  In some cases, this valve is affixed to the chamber, and in other cases, it’s provided as a loose item.

- The shutoff valve requires a ¼” o.d. supply line.  To connect this line to the valve, simply insert it into the valve body until it hits a positive stop.  The water valve’s push-type connection fitting will provide some resistance to the tubing before it properly seats.  Approximately a 5/8” length of the supply line tubing should insert into the valve when properly inserted.

## Fill the Humidifier with Water

Floor-mounted humidifiers utilize a mechanical float valve and should fill as soon as the supply valve is opened.

Evaporator-mounted humidifiers utilize a level switch and solenoid to achieve and maintain the proper water level inside the humidifier pan.  The humidifier will not fill until chamber power is turned on, and there are no active alarm states.

Open the water supply shut-off valve to fill the humidifier’s reservoir pan.  In its open position, the handle should be parallel to the valve body and water line.  Conversely, in its closed position the hand should be perpendicular to the valve body and water line.

Please note that the heater element(s) inside the humidifier water pan must be completely submerged with water to ensure proper operation.  If operated dry, the heaters can become excessively hot and damage the humidifier assembly and adjacent chamber componentry.

The water level in an evaporator-mounted humidifier assembly does not need to be verified prior to use as its float is factory-set, ships fully installed and has a low likelihood of falling out of adjustment.

The water level in a floor-mounted humidifier assembly should be verified prior to operation due to its installation requirements.  Remove the reservoir cover and observe the system as it fills with water.  The reservoir should fill with enough water to completely cover the heater element(s) before the valve closes.

## Float Valve Adjustment

The water fill level in floor-mounted humidifier reservoirs can be changed by adjusting the mechanical float valve position.  To make this adjustment:

- Remove the humidifier cover

- Loosen the thumb screw located between the float bulb and valve body

- Adjust the position of the float bulb, relative to the valve body

  - For a lower water level, adjust the position of the bulb to be higher

  - For a higher water level, adjust the position of the bulb to be lower

- Tighten the thumbscrew once the float bulb is in the desired position

- Place the humidifier cover back onto the reservoir

## Installation of Dehumidification Components

All provided dehumidification components described above are pre-installed and fully integrated into every chamber model at Percival’s factory prior to shipment.  No installation is required by the end user.

# Operation

## General

Relative humidity set points are achieved through a closed-loop control scenario.  The chamber’s control system continuously measures the relative humidity process value inside the controlled environment.  It compares the process value input from the sensor to the prescribed set point and adjusts operation of the humidifier and dehumidifier, depending on the difference between the values.

If the process value is lower than the set point:

- The humidifier assembly’s water pan heater(s) are energized

- The dehumidification system’s refrigeration solenoid valve(s) and electric air heater(s) are de-energized

If the process value is higher than the set point:

- The humidifier assembly’s water pan heater(s) are de-energized

- The dehumidification system’s refrigeration solenoid valve(s) and electric air heater(s) are energized

The on/off ratio of the %RH control system outputs are determined by independent PID control loops, which help to balance system responsiveness with control accuracy.  For example, if the process value is significantly lower than the set point, the control system will energize the humidifier for a large percentage of the cycle time.  During this time, the dehumidification system will be de-energized for most of its cycle time.  As the process value increases, approaching the set point, the system will begin to reduce the amount of time the humidifier is energized and increase the amount of time the dehumidifier is energized to prevent overshoot.

## Additive Humidification

The H3/H6 (Water Pan Humidifier with Dehumidification) system raises the %RH level inside Percival controlled environmental chambers via evaporation.  As air flows over the humidifier, it picks up moisture from the water mass.  Reservoir water temperature is increased via the submerged heater element(s), to increase the evaporation rate.

## Dehumidification

The H3/H6 system reduces relative humidity inside a chamber environment by condensing moisture out of the air onto cold, bare-tube evaporator coils.

Process air is passed through the main evaporator(s) for temperature conditioning, then passes through the dehumidification evaporator coil(s).  When enabled, the dehumidification evaporator operates at just above the freezing point.  As moisture condenses upon the cold coil surface, it sloughs off into a drip pan and is routed to the chamber floor drain.

Upon exiting the dehumidification evaporator(s), process air is passed over electric air heaters, where heat lost from the dehumidification process is reintroduced.

When the control system calls for heating and dehumidification, refrigerant only flows through the dehumidification evaporator(s).  Heating is accomplished solely through the chamber’s main electric air heaters during this condition.

To operate the humidity control system:

- Ensure the humidifier is installed and filled with water per the previous installation steps

- Enable Channels A and B of the %RH control loop in the Inputs Menu of the control system.  Please refer to the control system user manual for specific instruction on this process

- Set the desired relative humidity level via a Manual Set Point, or Multi-step Program.  Please refer to the control system user manual for specific instructions on this process

The %RH level inside the environment will self-adjust to the desired set point, provided the operational conditions described earlier in this manual are all satisfied.

Due to the psychrometric conditions within the controlled environment, achievement of the set point may take multiple hours.

## Relative Humidity Control Envelope

The chart below describes the approximate performance envelope for relative humidity control in most Percival controlled environment chambers.  (Per the performance specification section above, 22 and 30 series chambers are capable of dehumidification down to 50%RH).

As described in the chart above, it is possible to prescribe temperature and relative humidity parameter pairs that fall outside of the performance specification envelope.  Within the orange region of the performance envelope, achievement of the set point and control stability cannot be guaranteed.  There also exists an increased likelihood of excessive condensation, loss of process control, reduced part lifespan, and chamber malfunction within this region.

Please note that the lower %RH boundaries of the chart above are only achievable with optional dehumidification systems.

As a protective measure, the chamber’s control system automatically disables the humidification system if the set point is less than or equal to 10°C or greater than or equal to 40°C.

# Maintenance

## Clean the Floor-mounted System

- Shut off supply water to the chamber.

- Unplug the humidifier.

- Remove water from the humidifier via floor drain or a wet/dry shop-vac.

- Remove any debris that may have fallen or migrated into the pan.

- Clean all surfaces of the reservoir and lid with water and a mild detergent, or a solution of water and vinegar or other sanitizing agent as needed.  If needed, the humidifier assembly can be removed from the chamber for cleaning.  Please refer to installation instructions above for steps on removal/installation of the humidifier.

- Inspect the heater element(s) for corrosion, cracks or chips, and loose wiring during the cleaning process.

- Re-install the humidifier pan, connecting the water feed line and power cord, per guidelines from the installation section of this manual.

- Open the water supply valve.

Recommended Cleaning Schedule: as needed between experiments, or every 6 months

Recommendations for Storage: Completely drain water from the humidifier prior to any periods where the chamber will not be in use

## Flush Evaporator-mounted System

- Shut off supply water to the chamber.

- Drain water from the reservoir by removing the brass cap from the drain fitting, which is located on the front of the reservoir.  Some means to capture water as it drains will be necessary to prevent water from spilling into the chamber.  A drain hose, catch pan, or wet/dry shop-vac are all effective ways to drain the system.

- Allow the reservoir to fully drain.

- Re-attach the drain cap.

- Turn on supply water to allow the system to refill.

Recommended Cleaning Schedule: as needed between experiments, or every 3 months

Recommendations for Storage: Completely drain water from the humidifier prior to any periods where the chamber will not be in use

## Clean the Dehumidification Evaporator Coils

Dehumidification evaporator coil(s) are integrated into the main evaporator assembly and constructed of the same materials.  Inspect and clean the dehumidification coils in the same manner and schedule as the main evaporator, as described in the base chamber user manual.

# Troubleshooting

| Problem | Reason | Solution |
| --- | --- | --- |
| Reservoir won’t fill | Interrupted water supply | Check supply line for blockages and valve position |
|  | Non-opening fill valve | Check position of float bulb Clear blockage at valve inlet Check level switch and solenoid valve electrical system |
| Low reservoir water level | Float valve is out of adjustment | Check position of the float bulb and adjust as needed Check inlet valve for blockage, preventing it from closing |
| High reservoir water level |  |  |
| Heater element is not energizing | Electrical issue | Ensure heater element is fully seated into the appropriate receptacle Check the humidity system’s circuit breaker in the control box |
| Mineral deposits on chamber surfaces | Water quality | Check supply water quality against system requirements, specified above |
| %RH level won’t go below the ambient level | Specification limitations | Per specifications, the H1 system is additive only and cannot dehumidify |
| The relative humidity level won’t increase inside the chamber | Control system setup | Make sure Channel A is enabled in the RH inputs menu Set points may be outside of allowable operating range |
|  | Electrical | Check for tripped circuit breaker(s) in the control panel Make sure power cords are connected to appropriate receptacle and are fully seated |
|  | Mechanical | Check humidifier water level Open adjustable lid on floor-mounted system’s cover. Or remove lid |
|  | Environmental (very dry ambient) | Close fresh air exchange ports Limit door openings |

# Parts and Service

Replacement port plugs and caps can be ordered from Percival Scientific at:

Ph: 1-800-695-2743

Email: parts@percival-scientific.com

For additional questions regarding this product, please contact Percival’s customer support team at:

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com