IW - IncuWhite™ LED Lighting

Operator’s Manual

505 Research Drive  Perry, Iowa 50220

Toll-free 800.695.2743  515.465.9363  fax 515.465.9464

www.percival-scientific.com

[This page intentionally left blank]

# Introduction

To help customers create better science, Percival Scientific continuously studies and refines the design of its products to provide the best possible solutions and equipment in the industry.  As such, changes in the design and/or operation of equipment may occur without publication.  Please feel free to contact Percival Scientific directly for product updates.

Information contained within this supplemental user manual shall supersede any conflicting information contained within the basic chamber user manual.

Please read this manual carefully before proceeding with operation

# Revision Log

| Version | Date | Description |
| --- | --- | --- |
| Rev A | 8/15/2024 | Initial Release |

Introduction	3

Revision Log	4

Warranty Information	6

Definitions	8

General	8

Construction	8

Mounting	9

Technical Specifications	9

24 in IncuWhite™	9

48 in IncuWhite™	9

IncuWhite™ spectral output:	10

Installation	10

Incubators (I Series)	10

Vertical Lighting	10

Horizontal Lighting	10

Operation	13

Maintenance	14

Troubleshooting	14

Parts and Service	15

# Warranty Information

Equipment Limited Warranty

Thank you for purchasing equipment from Percival Scientific, Inc.  Our goal is to provide every customer with an exceptional ownership experience.  We are committed to providing products and services of the highest quality, and your satisfaction is our top priority.  In support of this objective, we present the following warranty.

Warranty Terms

Percival Scientific, Inc. (hereinafter referred to as Percival) warrants that equipment manufactured by Percival shall be free from defect in material or workmanship, when installed, operated and maintained in accordance with instructions provided by Percival.

Warranty Coverage

- Coverage period shall be twenty-four (24) months from the date of delivery

- Percival’s obligation under this warranty is limited to the repair or replacement of defect at no charge for parts, shipping and labor

- Percival reserves the right to inspect the equipment after a warranty claim is made to determine the proper course of corrective action

Exclusions

- Parts or elements damaged from improper installation, including mechanical, electrical and plumbing services, IT infrastructure, inadequate room ventilation, improper water quality, and ambient conditions that are out of specification

- Parts or elements damaged from improper operation

- Parts or elements damaged from improper maintenance

- Parts or elements considered to be consumable, that outlive their normal expected terms of usefulness, such as: light bulbs, filters, desiccant material and CO2 scrubbing media

- Damage resulting from shipment to or relocation within the customer’s facility after equipment has been signed for and ownership is transferred to the customer

- Modification to equipment by the customer without Percival’s consent may void this warranty

- Repair work undertaken without prior consent of Percival may void this warranty

- Damage caused by accident, fire, Force Majeure or Acts of God

Additional Terms

- This warranty is in lieu of any other warranties, expressed or implied, including merchantability or fitness for a particular purpose

- Warranty terms described herein cannot be changed without the expressed, written consent from Percival

- The owner agrees that Percival’s sole liability with respect to defective parts or materials shall be set forth in this warranty, and any claims for incidental or consequential damages are expressly excluded

- This warranty is non-transferable and applies only to the original purchaser

- Repaired or replaced parts shall be warranted only for the unexpired portion of the original warranty, unless otherwise defined

- Parts replaced under warranty shall become the property of Percival

To Make a Warranty Claim

- Document the issue with detailed notes and photographs along with the equipment’s serial number

- Contact Percival’s customer support team at customerservice@percival-scientific.com, parts@percival-scientific.com or 1-800-695-2743

- Once your claim is received, our customer support team will review it and contact you with further instructions

# Definitions

LED – Light Emitting Diode

‘Lamp bank’, ‘Light fixture’, ‘Fixture’ and ‘Lights’ are used interchangeably in this manual to describe the IncuWhite™ LED lighting system.

Light output in this manual is measured in µmol/m²/s and published output data was collected at 6 in (15.2 cm) from the source in an operating environment of 25°C

# General

LED lighting has become increasingly popular in controlled environment chambers due to overall system-level performance gains that can be realized from their use.  Compared to conventional light sources (e.g. fluorescent, incandescent and HID), LEDs are more energy efficient, have higher durability, longer life, and offer improved controllability and tunability.  LEDs can be tailored to deliver the specific quantity and quality of light required for simulating diurnal photoperiods and promoting optimal growth conditions for plants, microorganisms, and cell cultures.

LEDs introduce less heat into a controlled environment, which reduces demand on the unit’s cooling system and facilitates the stable control of critical process parameters.  The net effect is improved experimental efficacy, greater reproducibility, and reduced operational costs.

The IW (IncuWhite™ LED Lighting) option provides an LED lighting system designed specifically for use in Percival Scientific controlled environment chambers with low to medium light intensity requirements, such as biological incubators.

## Construction

IncuWhite™ LEDs are offered in 24” (61 cm) and 48” (122 cm) length form factors.  Both lengths share a common design where the LED circuit board(s) are affixed to an aluminum extrusion.  This extrusion is used for strength as well as heat dissipation.  Plastic caps with rubber seals and a light dispersing protective cover provide additional protection up to an IP65 rating for the product.

## Mounting

The IncuWhite™ LED modules’ aluminum housing snaps into various retainer clips, allowing for vertical or horizontal orientation upon a flat surface or wire shelf.

# Technical Specifications

## 24 in IncuWhite™

Length: 24 in (61 cm)

Voltage: 24 VDC

Wattage: 12W (each)

Intensity: ~30 µmol/m²/s at 6in (15.2cm) from the source

CCT = 5700K

CRI = 90

Control Range: 5 – 100% output

## 48 in IncuWhite™

Length: 48in (122 cm)

Voltage: 24 VDC

Wattage: 24W (each)

Intensity: ~60 µmol/m²/s at 6in (15.2cm) from the source

CCT = 5700K

CRI = 90

Control Range: 5 – 100% output

## IncuWhite™ spectral output:

# Installation

## Incubators (I Series)

IncuWhite™ LEDs are the standard lighting system utilized in Percival Scientific’s line of biological incubators.

### Vertical Lighting

Vertical lighting (included on all incubator model numbers with a ‘VL’ designation) ships pre-installed from Percival’s factory and requires no installation by the User.

### Horizontal Lighting

Horizontal lighting (included on all incubator model numbers with an ‘LL’ designation) is affixed to the underside of the unit’s wire shelving, which is shipped uninstalled from Percival’s factory.  To install horizontal lights on ‘LL’ models:

- Unpackage each shelf and locate shelf hardware (pilaster clips)

- Install pilaster clips:

  - Four clips, located at the same height are required to properly support each shelf

  - Numbers are embossed into each pilaster to assist with alignment of the clips

  - Insert the upper pilaster clip tab into the pilaster first

  - Twist the pilaster clip in a downward motion and insert the lower tab into the lower pilaster slot.

Note:  It may be necessary to pinch the clip slightly to obtain proper engagement of both clip tabs into the pilaster.

The image below shows an example of a properly installed pilaster clip:

- Place shelves onto the installed pilaster clips, ensuring that the vertical tab on each clip sits inboard of the shelf frame and secures each shelf from sliding off any clip.  When properly installed a shelf can only be removed by lifting it vertically off its clips.

The image below shows an example of a properly installed wire shelf:

- If a shelf has pre-installed IncuWhite™ LED modules, make sure to orient the assembly such that power cords extend toward the back wall of the controlled environment, or are close in proximity with the chamber’s corresponding receptacle cords.

The image below shows an example of proper shelf and light orientation:

Ensure chamber power is turned off, prior to plugging in LEDs

- Insert the male power plug of each IncuWhite™ LED module into its corresponding female receptacle.  The connector pair is keyed such that they can only be connected in one orientation.

The images below show the plug and receptacle pair, and their proper, full engagement:

# Operation

IncuWhite™ LEDs are controlled through open-loop dimmable light outputs via the chamber’s IntellusUltra control system.  This control scenario allows the user to set lighting output level(s) as a percentage of total output.  For example, a set point of 100% will energize the associated LED modules to their maximum output and a set point of 50% will reduce light output to half of their maximum output.

IncuWhite™ LEDs are dimmable in 1% increments between 5 – 100% output.

Any set point less than 5% will de-energize the light output and turn off the associated lighting circuit(s).

The factory default setting for any IncuWhite lighting system is Off.

To operate the IncuWhite™ LEDs:

- Ensure the lights are installed according to the guidelines outlined above

- Set the desired light output level as either a Manual or Programmed set point

Please refer to the control system user manual for specific instructions on this process.

Upon entry of any new light output setting(s), the chamber’s control system sends a command signal to the associated light circuit(s), causing their power output to change.

The update frequency between the control system and downstream LED control devices is once per minute, or upon any saved setting change through the controller.  In certain circumstances, light output may not immediately correspond to controller settings after a power cycle for up to one minute.

Please note that whenever any light is turned on through the control system, Day Calibration Offsets are applied to the Temperature and Relative Humidity (if equipped) process control loops.  During the transition between lights on to off (or vice versa), some fluctuations in process control are possible.

# Maintenance

IncuWhite LED modules require minimal maintenance.  Inspect and clean lights in the same manner and on the same schedule as described in the base chamber user manual.

The IntellusUltra control system keeps track of the total hours the IncuWhite™ LED lighting system is in operation over time and notifies the user when the lighting system has reached its average rated life (ARL).  When the controller’s Light Lifetime Alarm threshold has been reached, it is recommended that the system’s light output be measured with a quantum sensor and checked against the chamber’s original light intensity specification(s).

Please refer to the IntellusUltra control system user manual for specific instructions regarding Light Lifetime Alarms

# Troubleshooting

| Problem | Reason | Solution |
| --- | --- | --- |
| Lights don’t turn on | Problem with power cords | Check to make sure LED modules are plugged in properly and plug/receptacle pairs are fully seated |
|  | Electrical issue – tripped circuit breaker | Check status of circuit breaker in control box and reset as needed |
|  | Electrical Issue – power circuit (relay(s), LED driver(s) | Contact customer support |
| Lights don’t turn off | Control Issue – disconnected, or incorrect control signal | Cycle chamber power. If problem persists, contact customer support |
| Lights don’t dim | LED driver malfunction | Cycle chamber power. If problem persists, contact customer support |

# Parts and Service

Replacement parts can be ordered from Percival Scientific at:

Ph: 1-800-695-2743

Email: parts@percival-scientific.com

For additional questions regarding this product, please contact Percival’s customer support team at:

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com

Please have your chamber serial number available for faster and more efficient customer service.