PC – PetriClear® LED Lighting

Operator’s Manual

505 Research Drive  Perry, Iowa 50220

Toll-free 800.695.2743  515.465.9363  fax 515.465.9464

www.percival-scientific.com

[This page intentionally left blank]

# Introduction

To help customers create better science, Percival Scientific continuously studies and refines the design of its products to provide the best possible solutions and equipment in the industry.  As such, changes in the design and/or operation of equipment may occur without publication.  Please feel free to contact Percival Scientific directly for product updates.

Information contained within this supplemental user manual shall supersede any conflicting information contained within the basic chamber user manual.

Please read this manual carefully before proceeding with operation

# Revision Log

| Version | Date | Description |
| --- | --- | --- |
| Rev A | 2/19/2025 | Initial Release |
| Rev B | 1/13/2026 | Changed low dimming range for 22, 30 and 36 to 1% from 5% |

Introduction	3

Revision Log	4

Warranty Information	6

Definitions	8

General	8

Construction	8

Mounting of Individual Panels	9

LED Panel Removal and Reinstallation	9

Technical Specifications	11

White LED Chip Color	11

PC - PetriClear® (12 x 24) Panel	11

PetriClear® Lamp Banks	13

Installation	13

Operation	16

White LED Circuit	16

Infrared LED Circuit	17

Maintenance	17

Troubleshooting	18

Parts and Service	19

# Warranty Information

Equipment Limited Warranty

Thank you for purchasing equipment from Percival Scientific, Inc.  Our goal is to provide every customer with an exceptional ownership experience.  We are committed to providing products and services of the highest quality, and your satisfaction is our top priority.  In support of this objective, we present the following warranty.

Warranty Terms

Percival Scientific, Inc. (hereinafter referred to as Percival) warrants that equipment manufactured by Percival shall be free from defect in material or workmanship, when installed, operated and maintained in accordance with instructions provided by Percival.

Warranty Coverage

- Coverage period shall be twenty-four (24) months from the date of delivery

- Percival’s obligation under this warranty is limited to the repair or replacement of defect at no charge for parts, shipping and labor

- Percival reserves the right to inspect the equipment after a warranty claim is made to determine the proper course of corrective action

Exclusions

- Parts or elements damaged from improper installation, including mechanical, electrical and plumbing services, IT infrastructure, inadequate room ventilation, improper water quality, and ambient conditions that are out of specification

- Parts or elements damaged from improper operation

- Parts or elements damaged from improper maintenance

- Parts or elements considered to be consumable, that outlive their normal expected terms of usefulness, such as: light bulbs, filters, desiccant material and CO2 scrubbing media

- Damage resulting from shipment to or relocation within the customer’s facility after equipment has been signed for and ownership is transferred to the customer

- Modification of equipment by the customer without Percival’s consent may void this warranty

- Repair work undertaken without prior consent of Percival may void this warranty

- Damage caused by accident, fire, Force Majeure or Acts of God

Additional Terms

- This warranty is in lieu of any other warranties, expressed or implied, including merchantability or fitness for a particular purpose

- Warranty terms described herein cannot be changed without the expressed, written consent from Percival

- The owner agrees that Percival’s sole liability with respect to defective parts or materials shall be set forth in this warranty, and any claims for incidental or consequential damages are expressly excluded

- This warranty is non-transferable and applies only to the original purchaser

- Repaired or replaced parts shall be warranted only for the unexpired portion of the original warranty, unless otherwise defined

- Parts replaced under warranty shall become the property of Percival

To Make a Warranty Claim

- Document the issue with detailed notes and photographs along with the equipment’s serial number

- Contact Percival’s customer support team at customerservice@percival-scientific.com, parts@percival-scientific.com or 1-800-695-2743

- Once your claim is received, our customer support team will review it and contact you with further instructions

# Definitions

LED – Light Emitting Diode

‘Lamp bank’, Light fixture’, ‘Fixture’, ‘Panel’, ‘Tile’, and ‘Light’ are used interchangeably in this manual to describe the SciWhite® Series LED lighting system.

Light output in this manual is measured in µmol/m²/s and published output data was collected at 6 in (15.2 cm) from the source in an operating environment of 25°C

# General

LED lighting has become increasingly popular in controlled environment chambers due to overall system-wide performance gains that can be realized from their use.  Compared to conventional light sources (e.g. fluorescent, incandescent and HID), LEDs are more energy efficient, have higher durability, longer life, and offer improved controllability and tunability.  LEDs can be tailored to deliver the specific quantity and quality of light required for simulating diurnal photoperiods and promoting optimal growth conditions for plants, microorganisms, and cell cultures.

LEDs introduce less heat into a controlled environment, which reduces demand on the unit’s cooling system and facilitates the stable control of critical process parameters.  The net effect is improved experimental efficacy, greater reproducibility, and reduced operational costs.

PetriClear® LED modules are a patent-pending lighting system designed specifically for use in Percival Scientific’s CU Series tissue culture chambers.  The lighting system consists of general-purpose cool white LEDs that provide light in the visible spectrum, specifically PAR (400-700nm).  Specially tuned infrared LED chips are utilized to abate the formation of condensation inside petri dish lids.

## Construction

PetriClear® LEDs are offered in a 12 x 24 in nominal form factor and share a common design with SciWhite® Series hardware which consists of an aluminum substrate, a layer with etched circuits onto which LED chips are soldered, and a conformal coating which provides a protective outer layer.

## Mounting of Individual Panels

PetriClear® LED modules are integrated into lamp banks that are mounted horizontally above each tissue culture shelf.  Each LED panel must be affixed to a flat surface for proper thermal management.

Heat produced by a PetriClear® LED panel is dissipated through its aluminum substrate. Each panel should be installed such that the exposed aluminum side, opposite of the LED chips, is brought fully into contact with the mounting surface.

PetriClear® LED panels are pre-installed into chamber light fixtures at Percival’s manufacturing facility.  However, should the need arise, individual panels can be removed from a light fixture with a few basic tools:

- 1/8” (or smaller) flathead screwdriver

- #2 Philips screwdriver

### LED Panel Removal and Reinstallation

Ensure chamber power is turned off, and lamp banks are unplugged prior to any cleaning or service efforts.

- Remove electrical wiring from the LED panel’s connection terminals

  - Wiring is secured via spring clamps, located inside each wire terminal connection point

  - To open the spring clamp, apply pressure to the terminal with a small flat-head screwdriver as shown in the following image

  - Wiggle the wire free from its terminal while applying pressure with a 1/8” flat head screwdriver

- Remove all LED panel fasteners and shoulder washers with a Philips screwdriver

- Pull the LED panel free from its mounting surface, allowing for the freed electrical wires to pull through the wire pass-through hole and grommet (not shown in the photo above).

(SW 12 x 12 in panel shown for reference)

- To reinstall a PC panel onto its mounting surface, perform the disassembly steps above in reverse order.  It is very important to ensure that:

  - Wiring must be fully inserted into the properly marked terminal.  No exposed copper should be left visible at each connection point

  - Fasteners are self-threading and must not be over-tightened

# Technical Specifications

## White LED Chip Color

White LED chip output may vary according to the table below, resulting in perceptible differences in color from diode-to-diode and panel-to-panel.

| Nominal CCT (ANSI C78.377A) |  |
| --- | --- |
| CCT | CCT Range |
| 2700 K | 2725 K ± 145 |
| 3000 K | 3045 K ± 175 |
| 3500 K | 3465 K ± 245 |
| 4000 K | 3985 K ± 275 |
| 4500 K | 4503 K ± 243 |
| 5000 K | 5028 K ± 283 |
| 5700 K | 5665 K ± 355 |
| 6500 K | 6530 K ± 510 |

## PC - PetriClear® (12 x 24) Panel

Dimensions: 11.75 in (29.85 cm) x 24 in (60.96 cm)

Channel 1: Cool White

LED chip type: mid-power, cool white

CCT: 5700K

CIE 1931 (x,y): (0.330 ± 0.01, 0.345 ± 0.025)

LED Panel Voltage: 35.75V

Driver Type: Constant current

Control Range:

| Chamber Series | LED Control Range |
| --- | --- |
| CU-22 | 1 – 100% |
| CU-30 | 1 – 100% |
| CU-36 | 1 – 100% |
| CU-41 | 10 – 100% |

Spectral output:

| Color | Wavelength Range (nm) | Ratio (%) |
| --- | --- | --- |
| Blue | 400-500 | 25 |
| Green | 500-600 | 43 |
| Red | 600-700 | 29 |
| Far Red | 700-800 | 3 |

Spectrograph:

Channel 2: Infrared

LED chip type: low power, infrared

Wavelength peak: 860 nm

Half-wavelength width: 35 nm

LED Panel Voltage: 35.75V

Driver Type: Constant current

Control range: LEDs are constantly kept at 100% output

Spectrograph:

# PetriClear® Lamp Banks

PetriClear® LED panels are populated into lamp bank assemblies that fit into CU Series tissue culture chambers.  Lamp banks consist of a white, epoxy coated sheet metal housing which provides protection, structural support, wire management, heat dissipation and light reflection.  When properly installed, the lamp bank assemblies orient PetriClear® panels to properly illuminate each tier from above.

# Installation

All PetriClear® lighting ships pre-installed from Percival’s factory and requires no installation by the User.  Prior to first use:

- Remove and discard all cardboard shims and foam blocks used to secure lamp banks during shipping.

- Make sure lamp banks are not askew.  Each lamp bank should rest evenly onto its support rails and be fully inserted such that it comes into contact with the chamber’s interior rear panels.

The images below show an example of a properly installed above-shelf, horizontal SciWhite® Series lamp banks:

(CU-36L4 w/ PC Lighting)

Ensure chamber power is turned off, prior to plugging in lamp banks

- If unplugged, insert the male power plug of each lamp bank into its corresponding female receptacle.  The plug and receptacle are configured such that they can only be connected in one orientation.

Make sure each plug is fully seated into its receptacle with a firm press.  Failure to fully seat a receptacle into its receptacle may cause the lamp bank or chamber electrical system to malfunction or become damaged.

(Plug not fully seated)		             (Press plug firmly into receptacle)

# Operation

## White LED Circuit

PetriClear® White LEDs are controlled through open-loop dimmable light outputs via the chamber’s IntellusUltra control system.  This control scenario allows the user to set lighting output level(s) as a percentage of total output.  For example, a set point of 100% will energize the associated LED modules to their maximum output and a set point of 50% will reduce light output to half of their maximum output.

PetriClear® Series LEDs are dimmable in 1% increments.  Standard dimming range is dependent upon chamber base model.

The CU-41 series chamber model has a dimming range of 10 – 100%

All other CU series chamber models have a dimming range of 1 – 100%

Any set point less than the low dimming threshold percentage will de-energize the light output and turn off the associated white LED lighting circuit(s).

The factory default setting for PetriClear® white LED lighting is off.

To operate the PetriClear® LEDs:

- Ensure the lights are installed according to the guidelines outlined above

- Set the desired light output level as either a Manual or Programmed set point

Please refer to the control system user manual for specific instructions on this process.

Upon entry of any new light output setting(s), the chamber’s control system sends a command signal to the associated light circuit(s), causing their power output to change.

The update frequency between the control system and downstream LED control devices is once per minute, or upon any saved setting change through the controller.  In certain circumstances, light output may not immediately correspond to controller settings after a power cycle for up to one minute.

Please note that whenever any light is turned on through the control system, Day Calibration Offsets are applied to the Temperature and Relative Humidity (if equipped) process control loops.  During the transition between lights on to off (or vice versa), some fluctuations in process control are possible.

## Infrared LED Circuit

PetriClear® infrared LEDs are optimized to reduce condensation formation inside petri dish lids during both day (i.e. white LEDs energized) and night (i.e. white LEDs de-energized) periods and require no user input.  The infrared LED channel always remains energized, and no means of control for this circuit are provided to the User.

The factory default setting for PetriClear® infrared LED lighting is on.

Infrared LEDs are not visible by the human eye.  Operation of the infrared LED circuit cannot be assessed by visual inspection.

# Maintenance

PetriClear® LED modules require minimal maintenance.  If cleaning becomes necessary:

- Turn off chamber power

- Unplug the lamp bank

- Slide the lamp bank out of the chamber

- Inspect and clean the lamp bank sheet metal housing (avoid the LED panels) according to the general cleaning guidelines described in the base chamber user manual.

Prevent any liquid from getting inside the lamp bank housing through its vent holes.

- Gently clean LED panels only with a soft, dry microfiber cloth

- Once clean, plug in the lamp bank and chamber

The IntellusUltra control system keeps track of the total hours the PetriClear® white LEDs are in operation over time and notifies the user when the lighting system has reached its average rated life (ARL).  When the controller’s Light Lifetime Alarm threshold has been reached, it is recommended that the system’s light output be measured with a quantum sensor and checked against the chamber’s original light intensity specification(s).

Please refer to the IntellusUltra control system user manual for specific instructions regarding Light Lifetime Alarms

# Troubleshooting

| Problem | Reason | Solution |
| --- | --- | --- |
| Lights don’t turn on | Problem with power cords | Check to make sure LED modules are plugged in properly and plug/receptacle pairs are fully seated |
|  | Electrical issue – tripped circuit breaker | Check status of circuit breaker in control box and reset as needed |
|  | Electrical Issue – power circuit (relay(s), LED driver(s) | Contact customer support |
| Lights don’t turn off | Control Issue – disconnected, or incorrect control signal | Cycle chamber power. If problem persists, contact customer support |
| Lights don’t dim | LED driver malfunction | Cycle chamber power. If problem persists, contact customer support |

# Parts and Service

Replacement parts can be ordered from Percival Scientific at:

Ph: 1-800-695-2743

Email: parts@percival-scientific.com

For additional questions regarding this product, please contact Percival’s customer support team at:

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com

Please have your chamber serial number available for faster and more efficient customer service.