Q19 – External Drip Pan

Operator’s Manual

505 Research Drive  Perry, Iowa 50220

Toll-free 800.695.2743  515.465.9363  fax 515.465.9464

www.percival-scientific.com

[This page intentionally left blank]

# Introduction

To help customers create better science, Percival Scientific continuously studies and refines the design of its products in order to provide the best possible solutions and equipment in the industry.  As such, changes in the design and/or operation of equipment may occur without publication.  Please feel free to contact Percival Scientific directly for product updates.

Information contained within this supplemental user manual shall supersede any conflicting information contained within the basic chamber user manual.

Please read this manual carefully before proceeding with operation

# Revision Log

| Version | Date | Description |
| --- | --- | --- |
| Rev A | 6/26/2024 | Initial Release |

Introduction	3

Revision Log	4

Warranty Information	6

Definitions	8

General	8

Holding Capacity	8

Location	8

Installation	9

Operation	9

Maintenance	9

Troubleshooting	10

Fit	10

Condensation	10

Parts and Service	10

# Warranty Information

Equipment Limited Warranty

Thank you for purchasing equipment from Percival Scientific, Inc.  Our goal is to provide every customer with an exceptional ownership experience.  We are committed to providing products and services of the highest quality, and your satisfaction is our top priority.  In support of this objective, we present the following warranty.

Warranty Terms

Percival Scientific, Inc. (hereinafter referred to as Percival) warrants that equipment manufactured by Percival shall be free from defect in material or workmanship, when installed, operated and maintained in accordance with instructions provided by Percival.

Warranty Coverage

- Coverage period shall be twenty-four (24) months from the date of delivery

- Percival’s obligation under this warranty is limited to the repair or replacement of defect at no charge for parts, shipping and labor

- Percival reserves the right to inspect the equipment after a warranty claim is made in order to determine the proper course of corrective action

Exclusions

- Parts or elements damaged from improper installation, including mechanical, electrical and plumbing services, IT infrastructure, inadequate room ventilation, improper water quality, and ambient conditions that are out of specification

- Parts or elements damaged from improper operation

- Parts or elements damaged from improper maintenance

- Parts or elements considered to be consumable, that outlive their normal expected terms of usefulness, such as: light bulbs, filters, desiccant material and CO2 scrubbing media

- Damage resulting from shipment to or relocation within the customer’s facility after equipment has been signed for and ownership is transferred to the customer

- Modification to equipment by the customer without Percival’s consent may void this warranty

- Repair work undertaken without prior consent of Percival may void this warranty

- Damage caused by accident, fire, Force Majeure or Acts of God

Additional Terms

- This warranty is in lieu of any other warranties, expressed or implied, including merchantability or fitness for a particular purpose

- Warranty terms described herein cannot be changed without the expressed, written consent from Percival

- The owner agrees that Percival’s sole liability with respect to defective parts or materials shall be set forth in this warranty, and any claims for incidental or consequential damages are expressly excluded

- This warranty is non-transferable and applies only to the original purchaser

- Repaired or replaced parts shall be warranted only for the unexpired portion of the original warranty, unless otherwise defined

- Parts replaced under warranty shall become the property of Percival

To Make a Warranty Claim

- Document the issue with detailed notes and photographs along with the equipment’s serial number

- Contact Percival’s customer support team at customerservice@percival-scientific.com, parts@percival-scientific.com or 1-800-695-2743

- Once your claim is received, our customer support team will review it and contact you with further instructions

# Definitions

‘Pan’ and ‘Tray’ are used interchangeably in this manual to describe the drain pan included with the Q19 (External Drip Pan) optional equipment package.

# General

During a cooling cycle, the temperature control system in a controlled environment produces water condensation.  This condensate is typically viewed as an unwanted byproduct that must be managed and ultimately removed from the environment.

Most typically, waste condensate is captured in a drip tray located directly beneath the evaporator coil and is routed to the chamber’s floor drain via a gravity-fed drain line.

The Q19 (External Drip Pan) option provides a means by which rejected waste condensate from a Percival environmental chamber can be temporarily captured, if no drain facilities are available at the installation location.

This manual is intended to provide instructions on the installation, use/operation, maintenance and troubleshooting of the Q19 (External Drip Pan) option for Percival Scientific’s control environmental chambers.

## Holding Capacity

The drip tray can retain approximately 1 gallon of condensate.

## Location

The low-profile collection tray is designed to slide under the environmental chamber’s floor drain, external to the controlled environment.

# Installation

No tools are required for installation of this component.

The condensate drip tray is shipped inside the chamber, included with other chamber accessories and uninstalled parts.

To install the condensate tray, place it onto the facility floor and slide it under the environmental chamber.

Be sure to position the tray beneath the chamber’s floor drain, ensuring that the tray is at the lowest point of the chamber’s drainage system.

# Operation

To provide sufficient space for insertion of the condensate drip tray, the standard, external 90° elbow and 1” O.D. drain line have been removed from the chamber’s construction.

When placed under the chamber, beneath the floor drain stub, the drip tray can capture up to approximately 1 gallon (4 liters) of liquid.

The water level inside the drip tray should be monitored frequently to prevent overflow, and must be manually emptied periodically, when sufficiently full.

# Maintenance

Inspect the condensation tray during each use for damage that could compromise its ability to hold water.

Clean the condensation tray on the same schedule and in the same manner as described in the maintenance section of the chamber model’s operation manual.

# Troubleshooting

## Fit

It is very important to position the condensate drip tray directly beneath the chamber’s external floor drain stub.  Misalignment, or insufficient insertion under the chamber can lead to an uncontrolled water mess on the facility floor, and potential slip hazard.

## Condensation

Please note that the amount of condensate rejected from a controlled environment can fluctuate, and is highly dependent upon ambient conditions, interior environmental conditions, frequency of door openings, product load and optional humidity control systems.

# Parts and Service

Replacement parts can be ordered from Percival Scientific at:

Ph: 1-800-695-2743

Email: parts@percival-scientific.com

For additional questions regarding this product, please contact Percival’s customer support team at:

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com