Q30/Q31/Q32 – Carbon Dioxide Enrichment

Operator’s Manual

505 Research Drive  Perry, Iowa 50220

Toll-free 800.695.2743  515.465.9363  fax 515.465.9464

www.percival-scientific.com

[This page intentionally left blank]

# Introduction

To help customers create better science, Percival Scientific continuously studies and refines the design of its products to provide the best possible solutions and equipment in the industry.  As such, changes in the design and/or operation of equipment may occur without publication.  Please feel free to contact Percival Scientific directly for product updates.

Information contained within this supplemental user manual shall supersede any conflicting information contained within the basic chamber user manual.

Please read this manual carefully before proceeding with operation

# Revision Log

| Version | Date | Description |
| --- | --- | --- |
| Ver 0 | 9/8/25 | Draft |
| Rev A | 10/28/25 | Initial Release of Document |

Introduction	3

Revision Log	4

Warranty Information	6

Definitions	8

General	8

Performance Criteria	9

Performance Specifications	9

CO2 Sensor Range	9

CO2 Control Range	9

Sensing vs. Control	9

Process Value Stability	10

Installation Guidelines	10

Installation of Additive CO2 Components	10

CO2 Gas Supply Requirements	11

CO2 Gas Supply Line Installation	11

Operation	12

General	12

Additive CO2	13

CO2 Concentration Control Envelope	13

Exposure to Elevated CO2 Levels	14

OHSA Limits	15

CDC Limits	15

Control System Alarm Limits	15

Maintenance	15

Additive CO2 Systems	15

Troubleshooting	16

Parts and Service	17

# Warranty Information

Equipment Limited Warranty

Thank you for purchasing equipment from Percival Scientific, Inc.  Our goal is to provide every customer with an exceptional ownership experience.  We are committed to providing products and services of the highest quality, and your satisfaction is our top priority.  In support of this objective, we present the following warranty.

Warranty Terms

Percival Scientific, Inc. (hereinafter referred to as Percival) warrants that equipment manufactured by Percival shall be free from defect in material or workmanship, when installed, operated and maintained in accordance with instructions provided by Percival.

Warranty Coverage

- Coverage period shall be twenty-four (24) months from the date of delivery

- Percival’s obligation under this warranty is limited to the repair or replacement of defect at no charge for parts, shipping and labor

- Percival reserves the right to inspect the equipment after a warranty claim is made to determine the proper course of corrective action

Exclusions

- Parts or elements damaged from improper installation, including mechanical, electrical and plumbing services, IT infrastructure, inadequate room ventilation, improper water quality, and ambient conditions that are out of specification

- Parts or elements damaged from improper operation

- Parts or elements damaged from improper maintenance

- Parts or elements considered to be consumable, that outlive their normal expected terms of usefulness, such as: light bulbs, filters, desiccant material and CO2 scrubbing media

- Damage resulting from shipment to or relocation within the customer’s facility after equipment has been signed for and ownership is transferred to the customer

- Modification of equipment by the customer without Percival’s consent may void this warranty

- Repair work undertaken without prior consent of Percival may void this warranty

- Damage caused by accident, fire, Force Majeure or Acts of God

Additional Terms

- This warranty is in lieu of any other warranties, expressed or implied, including merchantability or fitness for a particular purpose

- Warranty terms described herein cannot be changed without the expressed, written consent from Percival

- The owner agrees that Percival’s sole liability with respect to defective parts or materials shall be set forth in this warranty, and any claims for incidental or consequential damages are expressly excluded

- This warranty is non-transferable and applies only to the original purchaser

- Repaired or replaced parts shall be warranted only for the unexpired portion of the original warranty, unless otherwise defined

- Parts replaced under warranty shall become the property of Percival

To Make a Warranty Claim

- Document the issue with detailed notes and photographs along with the equipment’s serial number

- Contact Percival’s customer support team at customerservice@percival-scientific.com, parts@percival-scientific.com or 1-800-695-2743

- Once your claim is received, our customer support team will review it and contact you with further instructions

# Definitions

‘Carbon Dioxide’ and ‘CO2’ are used interchangeably in this manual.

CO2 concentration is expressed in terms of parts per million (ppm) or as a percent concentration (%) of air.

‘Controlled environment chamber’, ‘Chamber’ and ‘Unit’ are used interchangeably in this manual to describe the complete controlled environment chamber system manufactured by Percival Scientific, Inc.

# General

Carbon dioxide concentration is an important variable in many controlled environment applications, and its precise and stable control is essential in creating physiologically relevant environments for cells and organisms.

Tight control of CO₂ allows an environmental chamber system to minimize variability, protect sensitive biological processes, and create standardized conditions necessary for reliable research, clinical trials, and regulatory compliance.

The Q30, Q31, and Q32 (CO2 enrichment) options provide the capability to raise CO2 levels inside Percival Scientific environmental chambers.  The system increases CO2 levels by injecting (customer supplied) pressurized CO2 gas into the environment.

This manual is intended to provide instructions on the installation, operation, maintenance and troubleshooting of the Q30, Q31, and Q32 additive CO2 options for Percival Scientific’s control environmental chambers.

# Performance Criteria

Performance of the Q30, Q31, and Q32 CO2 enrichment systems are based upon the following criteria:

- Pressure of CO2 being supplied to the chamber.

- Ambient CO2 levels around the chamber.

- A single point of control in the location of the CO2 sensor (typically located near the evaporator intake, or at shelf level adjacent to an aspirated sensor box)

- Empty chamber (i.e. no product load)

- An understanding that door openings, changes to the chamber’s control system or physical settings can temporarily destabilize CO2 level, causing process values (momentary measurement of the amount CO2  in the chamber) to deviate from their specified tolerance range for control

# Performance Specifications

## CO2 Sensor Range

| Option Code | Sensor Range | Control Accuracy (Stable State) |
| --- | --- | --- |
| Q30 | 0 – 2,000 ppm | ± 125 ppm around CO2 set point |
| Q31 | 0 – 5,000 ppm | ± 250 ppm around CO2 set point |
| Q32 | 0 – 10 % | ± 0.1% ppm around CO2 set point |

## CO2 Control Range

| Option Code | Control Range |
| --- | --- |
| Q30 | Above ambient to 1,800 ppm |
| Q31 | Above ambient to 4,800 ppm |
| Q32 | Above ambient to 9.8% |

## Sensing vs. Control

There is an important distinction between sensor range and control range.  Sensor range refers to the span of values a sensor can accurately detect or measure, whereas control range defines the span of values over which a system can effectively regulate a process variable.

In the context of a Percival environmental chamber, the CO2 sensor’s range specification describes the system’s ability to measure CO2 levels.  The sensor range is not a specification of the chamber’s operational range.

Ideally, sensor range should be marginally wider than control range.  Ensuring that the sensor range fully covers, but is appropriately aligned with, the control range is essential for precise and reliable system performance.

Ambient CO2 levels can vary significantly based on urban vs. rural settings, geographic location, and time of year.  These external variables establish the lower limit of the control and range for Q30, Q31, and Q32 additive CO2 packages.

## Process Value Stability

Process value stability is achieved after approximately 30 minutes of reaching the CO2 set point.  After achieving stable operating conditions, 95% (2σ) of the CO2 process values, based on the conditions above, will fall within the specified tolerance ranges.

# Installation Guidelines

Q30, Q31, and Q32 optional equipment packages are provided with all hardware, electrical infrastructure and control system configurations required for integration with most controlled environment chambers from Percival.  Please note that the specific implementation of this option will differ according to the base chamber model into which the CO2 system is added.  Despite these differences, the general concept, function and control of the Q30, Q31, and Q32 system remains the same across all variations.

Proceed with all general chamber installation guidelines contained within the base chamber’s user manual before proceeding with the following CO2 installation guidelines.

## Installation of Additive CO2 Components

Failure to follow the instructions below may cause malfunction and/or damage to the CO2 system or chamber. Please contact Percival Scientific with any installation questions.

Solenoid valve(s) and needle valve(s) assemblies are shipped pre-installed on Percival chambers and require no installation.  The CO2 connection point and required supply pressure is labelled.

## CO2 Gas Supply Requirements

  - Customer supplied CO2 tanks should be capable of regulation down to 4 psig and have sufficient capacity to reach the set point and maintain the required supply while the CO2 system is in operation

  - Supply line with appropriate fitting for your chamber; either a push-type for a ¼” o.d. tube or a ¼” Male SAE

## CO2 Gas Supply Line Installation

It is highly recommended that the CO2 gas supply line be purged of any debris prior to connection with the chamber’s CO2 system.  Even very small particles in the supply line can prevent the solenoid valve from functioning properly.

To make the CO2 gas supply connection:

- Connect an appropriate CO2 gas supply line (as described above) to the provided needle valve labeled “CO2 Supply”

- The CO2 needle valve is set at Percival’s factory and should not need adjustment. Please contact service if there is a need to adjust this, as it affects other settings and it can be difficult to tune.  The ideal CO2 pressure supplied to the system is subject to chamber size, experiment type, and desired system response and controllability. Keep in mind that raising the supply pressure to the CO2 solenoid valve will increase the system’s responsiveness but will decrease overall controllability. Alternatively, lowering the CO2 supply pressure will have the opposite effect on the system.

- Should the need to adjust the valve be necessary.

  - Remove the clip ring

  - Slide the red ring up and turn the dial to adjust flow

  - Adjust flow to desired setting

  - Slide red ring down and reinsert clip.

Connect CO2 supply here (SAE fitting shown)

Install clip to lock valve position

Red ring shown in locked position

# Operation

## General

CO2 concentration set points are achieved through a closed-loop control scenario.  The chamber’s control system continuously measures the CO2 concentration (process value) inside the controlled environment.  This process value input from the sensor is compared to the prescribed set point.  The control system adjusts the operation of the additive CO2 system, depending on the difference between the values.

Additive CO2 is accomplished by opening a solenoid valve to inject CO2 gas into the environment.

If the process value is lower than the set point:

- The CO2 supply solenoid valve(s) is energized

If the process value is higher than the set point:

- The CO2 supply solenoid valve(s) is de-energized

- Because the chamber is not perfectly sealed, CO2 concentration inside the environment will passively normalize toward ambient conditions over time.

The on/off ratio of the CO2 control system outputs are determined by independent PID control loops, which help balance system responsiveness with control accuracy.  For example, if the process value is significantly lower than the set point, the control system will energize the CO2 solenoid valve for a large percentage of the cycle time.  As the process value increases, approaching the set point, the system will begin to reduce the amount of time the CO2 solenoid valve is energized.

## Additive CO2

To increase CO2 concentration, set the desired level inside the chamber via a Manual Set Point, or Multi-step Program.

The CO2 level inside the environment will self-adjust to the desired set point, provided the operational conditions described earlier in this manual are all satisfied.

Please refer to the control system user manual for specific instruction on this process

Due to many factors including the conditions within the controlled environment and ambient CO2 levels, achievement of the set point may take up to approximately an hour.

## CO2 Concentration Control Envelope

The upper and lower bounds of the control envelope were determined by Percival Scientific by considering many factors, including leakage between the chamber and ambient conditions, ambient CO2 concentration, and additive CO2 gas flow rate. Chamber CO2 concentration can also be affected by products and specimens placed inside the chamber.

Because the Q30, Q31, and Q32 optional equipment packages are additive only, they increase CO2 concentration inside the environment more aggressively than it can passively decrease.  Because the system cannot quickly counteract the increase of CO2, it is common for the process value to overshoot set points above ambient.  For set points near the upper bound of the sensor range, there exists the possibility of over saturating the CO2 sensor.

If the CO2 sensor is over-saturated (i.e. when the CO2 level exceeds the upper limit of the CO2 sensor’s range), the sensor returns an incorrect zero value to the controller.

Should this happen, the control system has been configured to identify this situation, disable the supply of CO2, and alert the user through an alarm notification.

It is possible to change the setpoints of the chamber outside of the performance envelope, however it is not advised by Percival Scientific. Exceeding the upper bound of the control envelope is likely to cause loss of process control as described above.

As a protective measure, the chamber’s control system automatically disables the additive CO2 system if the controller detects that the CO2 sensor has become oversaturated.

# Exposure to Elevated CO2 Levels

This equipment can be set to operate at elevated CO2 levels. Improper use or inadequate ventilation can result in hazardous CO₂ concentrations as outlined below. Percival Scientific, Inc. assumes no responsibility or liability for any injury, illness, or damage resulting from exposure to elevated CO₂ levels during operation, maintenance, malfunction, or misuse of this chamber.

It is the user’s responsibility to ensure that all facility CO₂ safety precautions are followed, including – but not limited to - proper ventilation, continuous gas monitoring, and compliance with all applicable safety regulations. By using this equipment, the user accepts full responsibility for safe operation and adherence to those guidelines.

## OHSA Limits

| CO2 Level | Exposure Limit |
| --- | --- |
| 5,000 ppm (0.5%) | 8 hours |
| 30,000 ppm (3%) | 5 minutes |

Reference: https://www.osha.gov/chemicaldata/183

## CDC Limits

| CO2 Level | Exposure Limit | Possible Exposure Effect |
| --- | --- | --- |
| 50,000 ppm (5%) | 30 minutes | Signs of intoxication |
| 70,000 ppm (7%) | 1-2 minutes | Unconsciousness |
| 100,000 ppm (10%) |  | Dangerous to Life |

Reference: https://www.cdc.gov/niosh/idlh/124389.html

## Control System Alarm Limits

In accordance with the CDC limits described above, chambers with Q32 (0-10% CO2 sensor) are factory set to activate a soft high alarm condition if the CO2 process value exceeds 7% (70,000 ppm).

Please refer to the control system user manual for specific instruction on this process

# Maintenance

## Additive CO2 Systems

- Periodically check the CO2 supply tank pressure and replace the tank when the supply pressure gets low.

- Periodically make sure all fittings in the system are tight

# Troubleshooting

| Problem | Reason | Solution |
| --- | --- | --- |
| The CO2 level won’t increase inside the chamber | Control system setup | Make sure Channel A is enabled in the CO2 inputs menu Check the controller for correct voltage output to relays: approximately 5VDC Set point may be outside of possible operating range |
|  | Electrical | Check for tripped circuit breaker(s) in the control panel Check relays for proper function Make sure all wiring is connected and fully seated |
|  | Mechanical | Check CO2 supply to chamber Check solenoid for proper function Check needle valve position Verify no leaks; all access ports and openings appropriately closed or open; closed to increase the level, open to decrease the level |
| Sensor Input Error Message | Control system setup | Ensure CO2 sensor is receiving power and connected to Intellus controller Check for an over-saturation condition (if so, turn off CO2) |
|  | Electrical | Make sure all wiring is connected and fully seated Check for tripped circuit breaker(s) in the control panel |
|  | Mechanical | Check solenoid valve for proper function – turn off CO2 gas supply if needed Check for oversaturation and reduce CO2 concentration inside environment |

# Parts and Service

Replacement port plugs and caps can be ordered from Percival Scientific at:

Ph: 1-800-695-2743

Email: parts@percival-scientific.com

For additional questions regarding this product, please contact Percival’s customer support team at:

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com