Q33 Carbon Dioxide Removal System

Operator’s Manual

505 Research Drive  Perry, Iowa 50220

Toll-free 800.695.2743  515.465.9363  fax 515.465.9464

www.percival-scientific.com

[This page intentionally left blank]

# Introduction

To help customers create better science, Percival Scientific continuously studies and refines the design of its products to provide the best possible solutions and equipment in the industry.  As such, changes in the design and/or operation of equipment may occur without publication.  Please feel free to contact Percival Scientific directly for product updates.

Information contained within this supplemental user manual shall supersede any conflicting information contained within the basic chamber user manual.

Please read this manual carefully before proceeding with operation

# Revision Log

| Version | Date | Description |
| --- | --- | --- |
| Ver 0 | 9/10/25 | Draft |
| Rev A | 10/28/25 | Initial document release |

Introduction	3

Revision Log	4

Warranty Information	6

Definitions	8

General	8

Performance Criteria	9

Performance Specifications	9

CO2 Sensor Range and Control Range	9

CO2 Level Control Range for Reach-In Chambers	9

CO2 Level Control Range for Walk-In Chambers	9

Process Value Stability	10

Installation Guidelines	10

Installation of CO2 Scrubbing Components	10

Operation	13

General	13

Additive CO2	13

CO2 Scrubbing	14

CO2 Concentration Control Envelope	14

Maintenance	15

CO2 Removal System	15

Troubleshooting	16

Parts and Service	17

# Warranty Information

Equipment Limited Warranty

Thank you for purchasing equipment from Percival Scientific, Inc.  Our goal is to provide every customer with an exceptional ownership experience.  We are committed to providing products and services of the highest quality, and your satisfaction is our top priority.  In support of this objective, we present the following warranty.

Warranty Terms

Percival Scientific, Inc. (hereinafter referred to as Percival) warrants that equipment manufactured by Percival shall be free from defect in material or workmanship, when installed, operated and maintained in accordance with instructions provided by Percival.

Warranty Coverage

- Coverage period shall be twenty-four (24) months from the date of delivery

- Percival’s obligation under this warranty is limited to the repair or replacement of defect at no charge for parts, shipping and labor

- Percival reserves the right to inspect the equipment after a warranty claim is made to determine the proper course of corrective action

Exclusions

- Parts or elements damaged from improper installation, including mechanical, electrical and plumbing services, IT infrastructure, inadequate room ventilation, improper water quality, and ambient conditions that are out of specification

- Parts or elements damaged from improper operation

- Parts or elements damaged from improper maintenance

- Parts or elements considered to be consumable, that outlive their normal expected terms of usefulness, such as: light bulbs, filters, desiccant material and CO2 scrubbing media

- Damage resulting from shipment to or relocation within the customer’s facility after equipment has been signed for and ownership is transferred to the customer

- Modification to equipment by the customer without Percival’s consent may void this warranty

- Repair work undertaken without prior consent of Percival may void this warranty

- Damage caused by accident, fire, Force Majeure or Acts of God

Additional Terms

- This warranty is in lieu of any other warranties, expressed or implied, including merchantability or fitness for a particular purpose

- Warranty terms described herein cannot be changed without the expressed, written consent from Percival

- The owner agrees that Percival’s sole liability with respect to defective parts or materials shall be set forth in this warranty, and any claims for incidental or consequential damages are expressly excluded

- This warranty is non-transferable and applies only to the original purchaser

- Repaired or replaced parts shall be warranted only for the unexpired portion of the original warranty, unless otherwise defined

- Parts replaced under warranty shall become the property of Percival

To Make a Warranty Claim

- Document the issue with detailed notes and photographs along with the equipment’s serial number

- Contact Percival’s customer support team at customerservice@percival-scientific.com, parts@percival-scientific.com or 1-800-695-2743

- Once your claim is received, our customer support team will review it and contact you with further instructions

# Definitions

‘Carbon Dioxide’ and ‘CO2’ are used interchangeably in this manual. CO2 concentration is expressed in terms of Parts Per Million (PPM) or as a percentage of the air.

# General

Carbon dioxide concentration is an important variable in many controlled environment applications, and its precise and stable control is essential in creating physiologically relevant environments for cells and organisms.

Tight control of CO₂ allows an environmental chamber system to minimize variability, protect sensitive biological processes, and create standardized conditions necessary for reliable research, clinical trials, and regulatory compliance.

Percival’s Q33 (CO2 removal) option provides the capability to actively lower CO2 levels inside Percival Scientific environmental chambers.  The system reduces CO2 levels via a dedicated CO2 scrubber and is sold as a complement to Q30, Q31, or Q32 (additive CO2) options.

This manual is intended to provide instructions on the installation, operation, maintenance and troubleshooting of the Q33 CO2 removal option for Percival Scientific’s control environmental chambers.

Additive CO2 options are discussed in a separate, dedicated user manual.

# Performance Criteria

Performance of the Q33 CO2 removal control system shall be based upon the following criteria:

- Amount, type and condition of CO2 scrubbing media

- Ambient CO2 levels around the chamber

- A single point of control in the location of the CO2 sensor (typically located near the evaporator intake, or at shelf level adjacent to an aspirated sensor box)

- Empty chamber (i.e. no product load)

- An understanding that door openings, changes to the chamber’s control system or physical settings can temporarily destabilize CO2 level, causing process values to deviate from their specified tolerance range for control

# Performance Specifications

## CO2 Sensor Range and Control Range

Please refer to the chamber equipment list and additive CO2 system user manual for details for your chamber’s CO2 sensor range and control specifications

## CO2 Level Control Range for Reach-In Chambers

| Option Code | Control Range |
| --- | --- |
| Q30/Q33 | 150 ppm to 1,800 ppm |
| Q31/Q33 | 150 ppm to 4,800 ppm |
| Q32/Q33 | 0.1% to 9.8% |

## CO2 Level Control Range for Walk-In Chambers

| Option Code | Control Range |
| --- | --- |
| Q30/Q33 | 250 ppm to 1,800 ppm |
| Q31/Q33 | 250 ppm to 4,800 ppm |
| Q32/Q33 | 0.1% to 9.8% |

## Process Value Stability

Process value stability is achieved after approximately 30 minutes of reaching the CO2 set point.  After achieving stable operating conditions, 95% of the CO2 process values, based on the conditions above, will fall within the specified tolerance ranges given.

Over time, the scrubbing system’s media will saturate with absorbed CO2, lose efficacy and eventually lose its ability to maintain below-ambient set points.  If the CO2 process value becomes unstable, or fails to achieve the set point, check the condition of the scrubbing media, and replace as necessary according to the step outlined later in this user manual.

# Installation Guidelines

The Q33 optional equipment package is provided with all hardware, electrical infrastructure and control system configurations required for integration with most controlled environment chambers from Percival.  Please note that the specific implementation of this option will differ according to the base chamber model into which the CO2 system is added.  Despite these differences, the general concept, function and control of the Q33 system remains the same across all variations.

Proceed with all general chamber installation guidelines contained within the base chamber’s user manual before proceeding with the following CO2 scrubber installation guidelines.

## Installation of CO2 Scrubbing Components

Failure to follow these instructions may cause malfunction or damage to the CO2 scrubbing system or chamber. Please contact Percival Scientific with any installation questions.

CO2 scrubber assemblies are pre-installed on Percival chambers and require no installation. The end user is responsible for filling the tray(s) with CO2 scrubbing material.

Preparing the System for CO2 Scrubbing

- To prepare the CO2 removal system for use, open CO2 scrubber’s main access door.

- Install the CO2 absorbent tray(s) as necessary. The trays rest on supports provided in the enclosure.

- Fill the tray(s) with CO2 absorbent.

  - Units are shipped with a 7.5 lb. bag of Sofnolime® (or similar) CO2-absorbing material.  End users are responsible for supplying additional scrubbing material

  - For proper air flow thru the material, all trays should be filled to a uniform depth of an approximately 1”

  - For maximum scrubbing effectiveness, and longest operation before replacement of scrubbing media becomes necessary, fill all provided trays as described above.

- Remove all port plugs installed between the CO2 scrubber and chamber to facilitate air exchange between the two spaces.

- Close the main access door, ensuring that a proper seal is established to prevent unwanted air leaks.  Gaskets should be compressed, and there should be no visible gaps between the access panel and scrubber enclosure.

- Enable Channel B in the CO2 inputs menu on the control system’s user interface

Please refer to the control system user manual for specific instruction on this process

Fill tray(s) with absorbent.  Empty trays may be removed as necessary.

Remove port plug when system is utilized.  Install plug in hole when the CO2 removal system is not being utilized.

Install CO2 absorbent tray(s)

Blower

Failure to disable channel B of the CO2 control on the Intellus controller when the port plug is installed may result in the failure of the blower from overheating.

# Operation

## General

CO2 concentration set points are achieved through a closed-loop control scenario.  The chamber’s control system continuously measures the CO2 concentration process value inside the controlled environment.  This process value input from the sensor is compared to the prescribed set point.  The control system adjusts the operation of the additive CO2 system and CO2 removal system, depending on the difference between the values.

Additive CO2 is accomplished by opening a solenoid valve to inject CO2 gas into the environment.

If the process value is lower than the set point:

- The CO2 supply solenoid valve(s) is energized

- The CO2 scrubber system air circulation blower(s) is de-energized

If the process value is higher than the set point:

- The CO2 supply solenoid valve(s) is de-energized

- The CO2 scrubber system air circulation blower(s) is energized

The on/off ratio of the CO2 control system outputs are determined by independent PID control loops, which help balance system responsiveness with control accuracy.  For example, if the process value is significantly lower than the set point, the control system will energize the CO2 solenoid valve for a large percentage of the cycle time.  During this time, the CO2 scrubber system will be de-energized for most of its cycle time.  As the process value increases, approaching the set point, the system will reduce the amount of time the CO2 solenoid valve is energized and increase the amount of time the CO2 scrubber system is energized to prevent overshoot.

## Additive CO2

Q30, Q31, or Q32 additive systems raise the CO2 concentration inside Percival controlled environmental chambers by injecting CO2 gas into the chamber’s process air stream.

Please refer to the additive CO2 system user manual for specific instructions on this process

## CO2 Scrubbing

The Q33 scrubber system reduces CO2 concentration inside a chamber environment by passing chamber air through scrubbing material, which reacts with CO2 molecules and removes them from the air.

Chamber air is brought into the CO2 scrubber’s enclosure by the scrubber’s blower(s) where it is then passed through scrubbing material. Once the air has passed through the scrubbing material, it exits the CO2 scrubber’s enclosure and returns to the chamber environment.

To operate the CO2 removal system:

- Ensure the CO2 supply and CO2 scrubber are installed according to their prescribed installation steps

- Enable Channels A and B of the CO2 control loop in the Inputs Menu of the control system.

- Set the desired CO2 concentration level via a Manual Set Point, or Multi-step Program.

Please refer to the control system user manual for specific instruction on this process

The CO2 level inside the environment will self-adjust to the desired set point, provided the operational conditions described earlier in this manual are all satisfied.

Due to many factors including the conditions within the controlled environment, ambient CO2 levels, and the condition of the scrubbing material, achievement of the set point may take multiple hours.

## CO2 Concentration Control Envelope

The upper and lower bounds of the control envelope were determined by Percival Scientific by considering many factors, including leakage between the chamber and ambient conditions, the ambient CO2 concentration, additive CO2 gas flow rate, volume of scrubbing material and volumetric flow rate of the scrubbing system’s blowers.

If CO2 concentration inside the chamber exceeds the upper limit of the sensing range of the CO2 sensor, the sensor will return an incorrect reading to the controller. Should this happen, the control system has been configured to identify this situation, disable the supply of CO2, and alert the user.

It is possible to change the setpoints of the chamber outside of the performance envelope, however it is not advised by Percival Scientific. Exceeding the upper bound of the control envelope as the control system will be unable to read the CO2 concentration, which will cause loss of process control. Exceeding the lower bound of the control envelope is also not advised as there are physical limitations to achieving extremely low CO2 levels and it may cause excessive wear on the CO2 scrubber’s fan(s).

As a protective measure, the chamber’s control system automatically disables the additive CO2 system if the controller detects that the CO2 sensor has become oversaturated.

# Maintenance

## CO2 Removal System

- Periodically open the CO2 scrubber compartment and inspect the scrubbing material. Based on the size of the chamber, type of experiment, ambient CO2 levels, and other variables, the replacement frequency of fresh scrubbing material will vary.  If there is a noticeable difference in chamber performance regarding CO2, check the scrubbing material. The material will change color when it needs to be changed.

  - The Sofnolime® scrubbing material supplied with the chamber will become discolored as it absorbs CO2, changing from white to violet. To replace used material, remove the storage tray(s), dump the old material, and replace it with fresh absorbent. Reinstall tray(s) with fresh absorbent in the scrubber compartment as described above.

  - Generally, used granules can be bagged and disposed of in ordinary landfill trash.

CO2 scrubbing media is considered to be an irritant as it contains calcium hydroxide and trace amounts of corrosive sodium hydroxide.

# Troubleshooting

| Problem | Reason | Solution |
| --- | --- | --- |
| The CO2 level won’t decrease inside the chamber | Control system setup | Ensure Channel B is enabled in the CO2 inputs menu Check the controller for correct voltage output to relays: approximately 5VDC Set point may be outside of possible operating range |
|  | Electrical | Check for tripped circuit breaker(s) in the control panel Check relays for proper function Make sure all wiring is connected and fully seated |
|  | Mechanical | Check CO2 supply pressure and needle valve position Verify supply solenoid is closed Check scrubbing material condition, replace as needed Check scrubber blower(s) for proper function and unimpeded airflow between CO2 scrubber and main chamber compartments Ensure exterior access ports and penetrations are closed, port plugs between scrubber and chamber are removed, and chamber door is properly closed |
| Sensor Input Error Message | Control system setup | Ensure CO2 sensor is receiving power and connected to Intellus controller Check for an over-saturation condition (if so, turn off CO2 control channel (B) |
|  | Electrical | Make sure all wiring is connected and fully seated Check for tripped circuit breaker(s) in the control panel |
|  | Mechanical | Check solenoid valve for proper function – turn off CO2 gas supply if needed Check for oversaturation and reduce CO2 concentration inside environment |

# Parts and Service

Replacement port plugs and caps can be ordered from Percival Scientific at:

Ph: 1-800-695-2743

Email: parts@percival-scientific.com

For additional questions regarding this product, please contact Percival’s customer support team at:

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com