SW – SciWhite® Series LED Lighting

Operator’s Manual

505 Research Drive  Perry, Iowa 50220

Toll-free 800.695.2743  515.465.9363  fax 515.465.9464

www.percival-scientific.com

[This page intentionally left blank]

# Introduction

To help customers create better science, Percival Scientific continuously studies and refines the design of its products to provide the best possible solutions and equipment in the industry.  As such, changes in the design and/or operation of equipment may occur without publication.  Please feel free to contact Percival Scientific directly for product updates.

Information contained within this supplemental user manual shall supersede any conflicting information contained within the basic chamber user manual.

Please read this manual carefully before proceeding with operation

# Revision Log

| Version | Date | Description |
| --- | --- | --- |
| Rev A | 2/04/2025 | Initial Release |

Introduction	3

Revision Log	4

Warranty Information	6

Definitions	8

General	8

Construction	8

SciWhite® Series Variants	9

Mounting of Individual Panels	9

Technical Specifications	11

White LED Chip Color	11

SW – SciWhite® (12 x 12) Panel	12

SW – SciWhite® (12 x 24) Panel	13

SWR – SciWhite® with Red (12 x 12) Panel	14

SWR – SciWhite® with Red (12 x 24) Panel	15

SWR+ – SciWhite® with Enhanced Red (12 x 12) Panel	16

SWR+ – SciWhite® with Enhanced Red (12 x 24) Panel	17

SW2 – SciWhite® with Independent WW and CW (12 x 24) Panel	18

SciWhite® Series Lamp Banks	21

Installation	21

Vertical Lighting (E-36VL and LT-36VL)	21

Fixed Canopy and Counterbalanced Lighting	21

Above Shelf, Horizontal Lighting	22

Operation	25

Maintenance	26

Troubleshooting	27

Parts and Service	28

# Warranty Information

Equipment Limited Warranty

Thank you for purchasing equipment from Percival Scientific, Inc.  Our goal is to provide every customer with an exceptional ownership experience.  We are committed to providing products and services of the highest quality, and your satisfaction is our top priority.  In support of this objective, we present the following warranty.

Warranty Terms

Percival Scientific, Inc. (hereinafter referred to as Percival) warrants that equipment manufactured by Percival shall be free from defect in material or workmanship, when installed, operated and maintained in accordance with instructions provided by Percival.

Warranty Coverage

- Coverage period shall be twenty-four (24) months from the date of delivery

- Percival’s obligation under this warranty is limited to the repair or replacement of defect at no charge for parts, shipping and labor

- Percival reserves the right to inspect the equipment after a warranty claim is made to determine the proper course of corrective action

Exclusions

- Parts or elements damaged from improper installation, including mechanical, electrical and plumbing services, IT infrastructure, inadequate room ventilation, improper water quality, and ambient conditions that are out of specification

- Parts or elements damaged from improper operation

- Parts or elements damaged from improper maintenance

- Parts or elements considered to be consumable or that outlive their normal expected terms of usefulness, such as: light bulbs, filters, desiccant material, and CO2 scrubbing media

- Damage resulting from shipment to or relocation within the customer’s facility after equipment has been signed for and ownership is transferred to the customer

- Modification to equipment by the customer without Percival’s consent may void this warranty

- Repair work undertaken without prior consent of Percival may void this warranty

- Damage caused by accident, fire, Force Majeure, or Acts of God

Additional Terms

- This warranty is in lieu of any other warranties, expressed or implied, including merchantability or fitness for a particular purpose

- Warranty terms described herein cannot be changed without the expressed, written consent from Percival

- The owner agrees that Percival’s sole liability with respect to defective parts or materials shall be set forth in this warranty, and any claims for incidental or consequential damages are expressly excluded

- This warranty is non-transferable and applies only to the original purchaser

- Repaired or replaced parts shall be warranted only for the unexpired portion of the original warranty, unless otherwise defined

- Parts replaced under warranty shall become the property of Percival

To Make a Warranty Claim

- Document the issue with detailed notes and photographs along with the equipment’s serial number

- Contact Percival’s customer support team at customerservice@percival-scientific.com, parts@percival-scientific.com or 1-800-695-2743

- Once your claim is received, our customer support team will review it and contact you with further instructions

# Definitions

LED – Light Emitting Diode

‘Lamp bank’, Light fixture’, ‘Fixture’, ‘Panel’, ‘Tile’ and ‘Light’ are used interchangeably in this manual to describe the SciWhite® Series LED lighting system.

Light output in this manual is measured in µmol/m²/s and published output data was collected at 6 in (15.2 cm) from the source in an operating environment of 25°C

# General

LED lighting has become increasingly popular in controlled environment chambers due to overall system-wide performance gains that can be realized from their use.  Compared to conventional light sources (e.g. fluorescent, incandescent, and HID), LEDs are more energy efficient, have higher durability, longer life, and offer improved controllability and tunability.  LEDs can be tailored to deliver the specific quantity and quality of light required for simulating diurnal photoperiods and promoting optimal growth conditions for plants, microorganisms, and cell cultures.

LEDs introduce less heat into a controlled environment, which reduces demand on the unit’s cooling system, and facilitates stable control of critical process parameters.  The net effect is improved experimental efficacy, greater reproducibility, and reduced operational costs.

SW (SciWhite® Series LED Lighting) options provide LED lighting systems designed specifically for use in Percival Scientific controlled environment chambers for a wide variety of applications.  General-purpose white LEDs are the foundational chip color upon which each SW panel is based.  Supplemental, wavelength-specific chips are strategically utilized to augment the output spectrum for specific controlled environment applications and product lines.

## Construction

SciWhite® Series LEDs are offered in 12 x 12 in and 12 x 24 in nominal form factors.  Both types share a common design consisting of an aluminum substrate, a layer with etched circuits onto which LED chips are soldered and a conformal coating which provides a protective outer layer.

## SciWhite® Series Variants

The SciWhite® Series LED platform includes multiple variants which are either provided as standard equipment, as described in each chamber’s base-model user manual, or can be ordered as optional equipment.  These include, but are not limited to:

| Option Code | Description | Notes |
| --- | --- | --- |
| SW | SciWhite® | 4000K white |
| SWR | SciWhite® with Red | 3500K white, 660nm red |
| SWR+ | SciWhite® with Enhanced Red | 3500K white, 660nm red |
| SW2 | SciWhite® with Independent CW and WW | 5700K white, 3000K white |

## Mounting of Individual Panels

All SciWhite® Series LED modules can be mounted in any orientation, provided there is sufficient flat surface area onto which each panel can be affixed.

Heat produced by SciWhite® Series LED panels is dissipated through the aluminum substrate layer. Each panel should be installed such that the exposed aluminum side, opposite of the LED chips, is brought fully into contact with the mounting surface.

Typically, SciWhite® LED panels are pre-installed into chamber light fixtures at Percival’s manufacturing facility.  However, should the need arise, individual panels can be removed from a light fixture with a few basic tools:

- 1/8” (or smaller) flathead screwdriver

- #2 Philips screwdriver

LED Panel Removal and Reinstallation

Ensure chamber power is turned off, and lamp banks are unplugged prior to any cleaning or service efforts.

- Remove electrical wiring from the LED panel’s connection terminals

  - Wiring is secured via spring clamps, located inside each wire terminal connection point

  - To open the spring clamp, apply pressure to the terminal with a small flat-head screwdriver as shown in the following image

  - Wiggle the wire free from its terminal while pressure is being applied with a 1/8” flat head screwdriver

- Remove all LED panel fasteners and shoulder washers with a Philips screwdriver

- Pull the LED panel free from its mounting surface, allowing for the freed electrical wires to pull through the wire pass-through hole and grommet (not shown in the photo above).

(SW 12 x 12 in panel shown for reference)

- To reinstall a SW panel onto its mounting surface, perform the disassembly steps above in reverse order.  It is very important to ensure that:

  - Wiring must be fully inserted into the properly marked terminal.  No exposed copper should be left visible at each connection point

  - Fasteners are self-threading and must not be over-tightened

  - All plastic bushings and shoulder washers are reinstalled

# Technical Specifications

## White LED Chip Color

White LED chip output may vary according to the table below, resulting in perceptible differences in color from diode-to-diode and panel-to-panel.

| Nominal CCT (ANSI C78.377A) |  |
| --- | --- |
| CCT | CCT Range |
| 2700 K | 2725 K ± 145 |
| 3000 K | 3045 K ± 175 |
| 3500 K | 3465 K ± 245 |
| 4000 K | 3985 K ± 275 |
| 4500 K | 4503 K ± 243 |
| 5000 K | 5028 K ± 283 |
| 5700 K | 5665 K ± 355 |
| 6500 K | 6530 K ± 510 |

## SW – SciWhite® (12 x 12) Panel

Dimensions: 11.75 in (29.85 cm) x 12 in (30.48 cm)

LED chip type: mid-power, warm white

PAR Efficacy: 2.165 µmol/J

CCT: 4000K

CIE 1931 (x,y): (0.380 ± 0.025, 0.381 ± 0.025)

Control Channel(s): 1

LED Panel Voltage: 35.75V

Driver Type: Constant current

Control Range: 10 – 100% output

Spectral output:

| Color | Wavelength Range (nm) | Ratio (%) |
| --- | --- | --- |
| Blue | 400-500 | 18 |
| Green | 500-600 | 39 |
| Red | 600-700 | 38 |
| Far Red | 700-800 | 5 |

Spectrograph:

## SW – SciWhite® (12 x 24) Panel

Dimensions: 11.75 in (29.85 cm) x 24 in (60.96 cm)

LED chip type: mid-power, warm white

PAR Efficacy: 2.380 µmol/J

CCT: 4000K

CIE 1931 (x,y): (0.380 ± 0.025, 0.381 ± 0.025)

Control Channel(s): 1

LED Panel Voltage: 35.75V

Driver Type: Constant current

Control Range: 10 – 100% output

Spectral output:

| Color | Wavelength Range (nm) | Ratio (%) |
| --- | --- | --- |
| Blue | 400-500 | 18 |
| Green | 500-600 | 39 |
| Red | 600-700 | 38 |
| Far Red | 700-800 | 5 |

Spectrograph:

## SWR – SciWhite® with Red (12 x 12) Panel

Dimensions: 11.75 in (29.85 cm) x 12 in (30.48 cm)

Primary LED chip type: mid-power, warm white

CCT: 3500K

CIE 1931 (x,y): (0.405 ± 0.025, 0.39 ± 0.025)

Supplemental LED chip type: mid-power, deep red

Wavelength: 660 nm

Control Channel(s): 1

LED Panel Voltage: 35.75V

Driver Type: Constant current

Control Range: 10 – 100% output

Spectral output:

| Color | Wavelength Range (nm) | Ratio (%) |
| --- | --- | --- |
| Blue | 400-500 | 14 |
| Green | 500-600 | 42 |
| Red | 600-700 | 41 |
| Far Red | 700-800 | 3 |

Spectrograph:

## SWR – SciWhite® with Red (12 x 24) Panel

Dimensions: 11.75 in (29.85 cm) x 24 in (60.96 cm)

Primary LED chip type: mid-power, warm white

CCT: 3500K

CIE 1931 (x,y): (0.405 ± 0.025, 0.39 ± 0.025)

Supplemental LED chip type: mid-power, deep red

Wavelength: 660 nm

Control Channel(s): 1

LED Panel Voltage: 35.75V

Driver Type: Constant current

Control Range: 10 – 100% output

Spectral output:

| Color | Wavelength Range (nm) | Ratio (%) |
| --- | --- | --- |
| Blue | 400-500 | 14 |
| Green | 500-600 | 42 |
| Red | 600-700 | 41 |
| Far Red | 700-800 | 3 |

Spectrograph:

## SWR+ – SciWhite® with Enhanced Red (12 x 12) Panel

Dimensions: 11.75 in (29.85 cm) x 12 in (30.48 cm)

Primary LED chip type: mid-power, warm white

CCT: 3500K

CIE 1931 (x,y): (0.405 ± 0.025, 0.39 ± 0.025)

Supplemental LED chip type: mid-power, deep red

Wavelength: 660 nm

Control Channel(s): 1

LED Panel Voltage: 35.75V

Driver Type: Constant current

Control Range: 10 – 100% output

Spectral output:

| Color | Wavelength Range (nm) | Ratio (%) |
| --- | --- | --- |
| Blue | 400-500 | 16 |
| Green | 500-600 | 36 |
| Red | 600-700 | 44 |
| Far Red | 700-800 | 4 |

Spectrograph:

## SWR+ – SciWhite® with Enhanced Red (12 x 24) Panel

Dimensions: 11.75 in (29.85 cm) x 24 in (60.96 cm)

Primary LED chip type: mid-power, warm white

CCT: 3500K

CIE 1931 (x,y): (0.405 ± 0.025, 0.39 ± 0.025)

Supplemental LED chip type: mid-power, deep red

Wavelength: 660 nm

Control Channel(s): 1

LED Panel Voltage: 35.75V

Driver Type: Constant current

Control Range: 10 – 100% output

Spectral output:

| Color | Wavelength Range (nm) | Ratio (%) |
| --- | --- | --- |
| Blue | 400-500 | 16 |
| Green | 500-600 | 36 |
| Red | 600-700 | 44 |
| Far Red | 700-800 | 4 |

Spectrograph:

## SW2 – SciWhite® with Independent WW and CW (12 x 24) Panel

Dimensions: 11.75 in (29.85 cm) x 24 in (60.96 cm)

Channel 1 LED chip type: mid-power, cool white

CCT: 5700K

CIE 1931 (x,y): (0.330 ± 0.01, 0.345 ± 0.025)

Control Channel: 1 of 2

LED Panel Voltage: 35.75V

Driver Type: Constant current

Control Range: 10 – 100% output

Spectral output:

| Color | Wavelength Range (nm) | Ratio (%) |
| --- | --- | --- |
| Blue | 400-500 | 25 |
| Green | 500-600 | 43 |
| Red | 600-700 | 29 |
| Far Red | 700-800 | 3 |

Spectrograph:

Channel 2 LED chip type: mid-power, warm white

CCT: 3000K

CIE 1931 (x,y): (0.430 ± 0.025, 0.405 ± 0.025)

Control Channel: 2 of 2

LED Panel Voltage: 35.75V

Driver Type: Constant current

Control Range: 10 – 100% output

Spectral output:

| Color | Wavelength Range (nm) | Ratio (%) |
| --- | --- | --- |
| Blue | 400-500 | 12 |
| Green | 500-600 | 36 |
| Red | 600-700 | 46 |
| Far Red | 700-800 | 6 |

Spectrograph:

Combined Output: Channels 1 (CW) and 2 (WW)

CCT: ~4500K

Spectral output:

| Color | Wavelength Range (nm) | Ratio (%) |
| --- | --- | --- |
| Blue | 400-500 | 12 |
| Green | 500-600 | 36 |
| Red | 600-700 | 46 |
| Far Red | 700-800 | 6 |

Spectrograph:

# SciWhite® Series Lamp Banks

SciWhite® Series LED panels are populated into lamp bank assemblies that integrate into a given chamber’s base model configuration.  Lamp banks consist of a white, epoxy coated sheet metal housing which provides protection, structural support, wire management, heat dissipation, and light reflection for SciWhite® panels.  Additionally, lamp bank assemblies orient SciWhite® panels for proper illumination (i.e. direction, intensity, and uniformity) of experiments inside Percival Scientific’s environmental chamber models.

The following table summarizes the standard implementation of SciWhite® lighting in Percival’s products:

| Chamber Series | SW Type | Orientation | Mounting Method |
| --- | --- | --- | --- |
| AL | SW | Above Shelf | Pilasters w/ clips |
| AR | SWR | Above Shelf | Pilasters w/ clips |
| E | SWR | Above Shelf | Pilasters w/ clips |
| E & LT-36/41VL | SWR | External, Vertical | Lift-off hinges |
| PGC-6/7 | SWR | Above Shelf | Pilasters w/ clips |
| PGC-9 | SWR | Above Shelf | Fixed |
| PGC-10/15/105 | SWR | Above Bench | Counterbalanced |
| PGC-20L1/40L2 | SWR | Above Shelf | Pilasters w/ clips |
| LT-105 | SWR | Above Bench | Counterbalanced |

# Installation

## Vertical Lighting (E-36VL and LT-36VL)

Vertical lighting (included on all reach-in plant growth chamber model numbers with a ‘VL’ designation) ships pre-installed from Percival’s factory and requires no installation by the User.

### Fixed Canopy and Counterbalanced Lighting

Fixed canopy and counterbalanced lighting systems ship pre-installed from Percival’s factory and requires no installation by the User.

### Above Shelf, Horizontal Lighting

After factory quality and performance testing is complete, horizontally mounted SW Series lamp bank assemblies are uninstalled, individually packaged and placed inside each chamber for shipping.

To install horizontal lamp banks:

- Unpackage each lamp bank and locate the mounting hardware (pilaster clips)

- Install pilaster clips:

  - Four clips, located at the same height are required to properly support each lamp bank

  - Numbers are embossed into each pilaster to assist with alignment of the clips

  - Insert the upper pilaster clip tab into the pilaster first

  - Twist the pilaster clip in a downward motion and insert the lower tab into the lower pilaster slot.

Note: It may be necessary to pinch the clip slightly to obtain proper engagement of both clip tabs into the pilaster.

The image below shows an example of a properly installed pilaster clip:

- Place lamp banks onto the installed pilaster clips, ensuring that the vertical tab on each clip sits inboard of the lamp bank frame and secures each assembly from sliding off any clip.  When properly installed a lamp bank can only be removed by lifting it vertically off its clips.

The images below shows an example of a properly installed above-shelf, horizontal SciWhite® Series lamp banks:

(AR-41L3 w/ SWR)

(Tabs on pilaster clips must sit inside lamp bank flanges)

- Make sure to orient each assembly such that power cords extend toward the back wall of the controlled environment or are close in proximity to the chamber’s corresponding receptacles.

The image below shows an example of a proper shelf and light orientation:

(AR-41 Series shown for reference)

Ensure chamber power is turned off, prior to plugging in lamp banks

- Insert the male power plug of each SciWhite® Series lamp bank into its corresponding female receptacle.  The plug and receptacle are configured such that they can only be connected in one orientation.

Make sure each plug is fully seated into its receptacle with a firm press.  Failure to fully seat a plug into its receptacle may cause the lamp bank or chamber electrical system to malfunction or become damaged.

(Plug not fully seated)		             (Press plug firmly into receptacle)

# Operation

SciWhite® Series LEDs are controlled through open-loop dimmable light outputs via the chamber’s IntellusUltra control system.  This control scenario allows the user to set lighting output level(s) as a percentage of total output.  For example, a set point of 100% will energize the associated LED modules to their maximum output and a set point of 50% will reduce light output to half of their maximum output.

SciWhite® Series LEDs are dimmable in 1% increments between 10 – 100% output.

Any set point less than 10% will de-energize the light output and turn off the associated lighting circuit(s).

The factory default setting for any SciWhite® lighting system is off.

To operate the SciWhite® Series LEDs:

- Ensure the lights are installed according to the guidelines outlined above

- Set the desired light output level as either a Manual or Programmed set point

Please refer to the control system user manual for specific instructions on this process.

Upon entry of any new light output setting(s), the chamber’s control system sends a command signal to the associated light circuit(s), causing their power output to change.

The update frequency between the control system and downstream LED control devices is once per minute, or upon any saved setting change through the controller.  In certain circumstances, light output may not immediately correspond to controller settings after a power cycle for up to one minute.

Please note that whenever any light is turned on through the control system, day calibration offsets are applied to the temperature and relative humidity (if equipped) process control loops.  During the transition between lights on to off (or vice versa), some fluctuations in process control are possible.

# Maintenance

SciWhite® Series LED modules require minimal maintenance.  If cleaning becomes necessary:

- Turn off chamber power

- Unplug the lamp bank

- Inspect and clean the lamp bank sheet metal housing (avoid the LED panels) according to the general cleaning guidelines described in the base chamber user manual.

Prevent any liquid from getting inside the lamp bank housing through its vent holes.

- Gently clean LED panels only with a soft, dry microfiber cloth

- Once clean, plug in the lamp bank and chamber

The IntellusUltra control system keeps track of the total hours the SciWhite® Series LED lighting system is in operation over time and notifies the user when the lighting system has reached its average rated life (ARL).  When the controller’s Light Lifetime Alarm threshold has been reached, it is recommended that the system’s light output be measured with a quantum sensor and checked against the chamber’s original light intensity specification(s).

Please refer to the IntellusUltra control system user manual for specific instructions regarding Light Lifetime Alarms

# Troubleshooting

| Problem | Reason | Solution |
| --- | --- | --- |
| Lights don’t turn on | Problem with power cords | Check to make sure LED modules are plugged in properly and plug/receptacle pairs are fully seated |
|  | Electrical issue – tripped circuit breaker | Check status of circuit breaker in control box and reset as needed |
|  | Electrical Issue – power circuit (relay(s), LED driver(s) | Contact customer support |
| Lights don’t turn off | Control Issue – disconnected, or incorrect control signal | Cycle chamber power. If problem persists, contact customer support |
| Lights don’t dim | LED driver malfunction | Cycle chamber power. If problem persists, contact customer support |

# Parts and Service

Replacement parts can be ordered from Percival Scientific at:

Ph: 1-800-695-2743

Email: parts@percival-scientific.com

For additional questions regarding this product, please contact Percival’s customer support team at:

Phone: 1-800-695-2743

Email: customerservice@percival-scientific.com

Please have your chamber serial number available for faster and more efficient customer service.